## Sheet1
Faraday shielded reference bridge. Rvb = 2K7 CF					Bridge evaluation by differential method. Version 1.00. DWK Oct. 2007.
		+/- esd					Cells with yellow background are input data.
R0cal =	49.90	0.06	<< Actual ref. load resistance (measured).
Design R0	50.00
k =	0.96	0.01
Ri =	50.04	0.06
N =	12
C2 + Cx	4.90		pF
Cx	0.00		pF	<< zero when Faraday shield is present.
C2	4.90
dR0 / dC1	0.817		Ohms / pF
x cal	5.231	0.005	<<< to determine the final calibration setting, vary xcal to get best true accuracy.
C1a cal	12.713		pF
				Magnitude error vs f
				f	C Scale	C1a	C1a – cal		R0	delta R0	|Z0|	delta|Z0|	|deltaZ0|
				/ MHz	x	/ pF	delta C1	delta R0	R0cal+dRo	/ R0		/ R0
				1.6	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006851
				1.7	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006856
				1.8	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006852
				1.9	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006853
				2.0	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006853
				2.2	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006858
				2.4	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006856
				2.6	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006857
				2.8	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006855
				3.0	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006851
				3.5	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006855
				4.0	5.270	12.582	-0.131	0.107	50.007	0.002139	50.007	0.000134	0.006722
				4.5	5.270	12.582	-0.131	0.107	50.007	0.002139	50.007	0.000134	0.006721
				5.0	5.270	12.582	-0.131	0.107	50.007	0.002139	50.007	0.000134	0.006721
				6.0	5.275	12.566	-0.147	0.120	50.020	0.002411	50.020	0.000406	0.020299
				7.0	5.275	12.566	-0.147	0.120	50.020	0.002411	50.020	0.000406	0.020298
				8.0	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006857
				9.0	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006858
				10.0	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006858
				12.0	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006858
				14.0	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006857
				16.0	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006854
				18.0	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006854
				20.0	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006853
				22.0	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006853
				24.0	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006854
				26.0	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006856
				28.0	5.265	12.599	-0.114	0.093	49.993	0.001867	49.993	-0.000137	0.006858
				30.0	5.260	12.616	-0.097	0.080	49.980	0.001594	49.980	-0.000409	0.020434
			Magnitude precision of the bridge +/- >>>					0.0204	Ohms
			Magnitude precision as a percentagee +/- >>>					0.0409	%
			True (claimable) accuracy of the bridge +/- >>>					0.0662	Ohms
			True accuracy as a percentagee +/- >>>					0.1325	%
KT =	11.575		<< Transformer constant.
C1 =	56.72		pF	<< estimated C1
C1b =	38.90		pF	<< padding capacitance. Nominal will do if no measurement is available.
C1s =	5.10		pF	<< estimated contribution to C1 from strays.

## Sheet2
Faraday shielded reference bridge. Rvb = 2K7 CF					T = 22 deg. C
				Cells with yelow background are input data.
R0	49.90
k	0.96
Ri	50.04
N	12
C2+Cx	4.90		pF	<<< must include strays.
Rvcal	2747.56		<<< to get final calibration setting, vary this to minimise max |phase error|
Li calc	8.062		uH	<<< should be ca. 1% lower than measured sec. Inductance. If not, C2+Cx has been wrongly estimated.
Lsec	8.150		uH	<<< measured sec. Inductance. For comparison
Li/Lsec	0.989

				f	Rv	Rv-Rvcal	dX0 / dRv	X0	delta phi	| delta phi |
				/MHz	/ Ohms		Deriv				Phase error vs f
				1.6	2750.0	2.4	0.010745	-0.026	-0.0300	0.0300
				1.7	2749.0	1.4	0.010120	-0.015	-0.0167	0.0167
				1.8	2750.0	2.4	0.009551	-0.023	-0.0267	0.0267
				1.9	2750.0	2.4	0.009048	-0.022	-0.0253	0.0253
				2.0	2750.0	2.4	0.008596	-0.021	-0.0240	0.0240
				2.2	2748.0	0.4	0.007826	-0.003	-0.0039	0.0039
				2.4	2746.0	-1.6	0.007184	0.011	0.0128	0.0128
				2.6	2746.0	-1.6	0.006632	0.010	0.0119	0.0119
				2.8	2745.0	-2.6	0.006162	0.016	0.0181	0.0181
				3.0	2743.0	-4.6	0.005760	0.026	0.0301	0.0301
				3.5	2744.0	-3.6	0.004933	0.018	0.0201	0.0201
				4.0	2744.0	-3.6	0.004317	0.015	0.0176	0.0176
				4.5	2745.0	-2.6	0.003834	0.010	0.0112	0.0112
				5.0	2744.0	-3.6	0.003453	0.012	0.0141	0.0141
				6.0	2744.0	-3.6	0.002878	0.010	0.0117	0.0117
				7.0	2750.0	2.4	0.002456	-0.006	-0.0069	0.0069
				8.0	2751.0	3.4	0.002147	-0.007	-0.0085	0.0085
				9.0	2749.0	1.4	0.001912	-0.003	-0.0032	0.0032
				10.0	2747.0	-0.6	0.001723	0.001	0.0011	0.0011
				12.0	2745.0	-2.6	0.001438	0.004	0.0042	0.0042
				14.0	2743.0	-4.6	0.001234	0.006	0.0065	0.0065
				16.0	2731.0	-16.6	0.001089	0.018	0.0207	0.0207
				18.0	2727.0	-20.6	0.000971	0.020	0.0229	0.0229
				20.0	2722.0	-25.6	0.000877	0.022	0.0257	0.0257
				22.0	2721.0	-26.6	0.000798	0.021	0.0243	0.0243
				24.0	2722.0	-25.6	0.000731	0.019	0.0214	0.0214
				26.0	2729.0	-18.6	0.000671	0.012	0.0143	0.0143
				28.0	2744.0	-3.6	0.000617	0.002	0.0025	0.0025
				30.0	2769.0	21.4	0.000565	-0.012	-0.0139	0.0139
				0.0301	Max |phase error|.

## Sheet3
Data reduction for 3-30pF beehive trimmer with reference markings..
		Dial	C /pF	esd C	fit	w	wy	wx	wx^2	wxy	y calc	obs-calc	we	we^2
		x	y					C calc	e
0 – 250pF range >>		0.0	30.0	0.25	1	16.000	480.00	0.00	0.00	0.00	30.10	-0.10	-1.61	0.163
		0.5	28.5	0.25	1	16.000	456.00	8.00	4.00	228.00	28.44	0.06	0.98	0.060
		1.0	26.5	0.25	1	16.000	424.00	16.00	16.00	424.00	26.78	-0.28	-4.43	1.224
		1.5	25.0	0.25	1	16.000	400.00	24.00	36.00	600.00	25.11	-0.11	-1.83	0.210
0-25pF range >>		2.0	23.7	0.10	1	100.000	2370.00	200.00	400.00	4740.00	23.45	0.25	24.76	6.129
		2.5	21.9	0.10	1	100.000	2190.00	250.00	625.00	5475.00	21.79	0.11	10.97	1.203
		3.0	20.1	0.10	1	100.000	2010.00	300.00	900.00	6030.00	20.13	-0.03	-2.82	0.080
		3.5	18.4	0.10	1	100.000	1840.00	350.00	1225.00	6440.00	18.47	-0.07	-6.62	0.438
		4.0	16.7	0.10	1	100.000	1670.00	400.00	1600.00	6680.00	16.80	-0.10	-10.41	1.083
		4.5	15.0	0.10	1	100.000	1500.00	450.00	2025.00	6750.00	15.14	-0.14	-14.20	2.016
		5.0	13.4	0.10	1	100.000	1340.00	500.00	2500.00	6700.00	13.48	-0.08	-7.99	0.638
		5.5	11.8	0.10	1	100.000	1180.00	550.00	3025.00	6490.00	11.82	-0.02	-1.78	0.032
		6.0	10.1	0.10	1	100.000	1010.00	600.00	3600.00	6060.00	10.16	-0.06	-5.57	0.310
		6.5	8.5	0.10	1	100.000	850.00	650.00	4225.00	5525.00	8.49	0.01	0.64	0.004
		7.0	6.9	0.10	1	100.000	690.00	700.00	4900.00	4830.00	6.83	0.07	6.85	0.469
		7.5	5.3	0.10	1	100.000	530.00	750.00	5625.00	3975.00	5.17	0.13	13.06	1.705
Starting to unmesh >>		8.0	4.1	0.10	0	0.000	0.00	0.00	0.00	0.00	3.51	0.59	0.00	0.000
		8.5	3.5	0.10	0	0.000	0.00	0.00	0.00	0.00	1.85	1.60	0.00	0.000
		9.0	3.1	0.10	0	0.000	0.00	0.00	0.00	0.00	0.18	2.92	0.00	0.000
		9.5	2.9	0.10	0	0.000	0.00	0.00	0.00	0.00	-1.48	4.38	0.00	0.000
		10.0	2.8	0.10	0	0.000	0.00	0.00	0.00	0.00	-3.14	5.94	0.00	0.000
		sums >>>>>>>>>			16	1264.000	18940.00	5748.00	30706.00	70947.00			8.32E-12	15.763
D	5772880
a	30.100796	Capacitance / pF vs logging scale
esd a	0.005989
b	-3.324183
esd b	0.000247
Chi^2/n-2)	1.125893
esd fit	1.061081