## Sheet1
FT50-61, 12 turns 0.9mm wire. Lsec = 9.25uH @ 1.5915MHz. 40pF variable ref cap with padding. Faraday shield earthed at gen. side.					D. W. Knight. June 2007.			Version	1.00
		esd +/-					Show columns N-Q to see deatails of fitting procedure.					See sheet 2 for C1'm vs frequency
N =	12.00					Li' = N R0 C2 Rv
R0 =	49.89	0.06					from fit	from fit
C2/pF	10.30	0.20					Ci =	Li =
L2 / nH	50.00	20.00					Li computed from Rv
L2 fit	1	< 1 to include L2 in fit, 0 to exclude it.
1/(N R0 C2)	1.62E+08					8.188	8.943
			Log (f)	f	x	Rv	esd Rv	Li'	Corr(L2)	y	esd y	fit	weight	w x	w x^2	w y	w x y	y calc	obs-calc	w e	w e^2	Li' calc	Li'					Computed	Computed
				/MHz	-(2 pi f)^2	Ohms	Ohms	/ uH					w					e			/ uH	obs-calc					Li /uH	Ci /pF
			6.15	1.4	-77.378	1462.4	14.62	9.0177	-4.4	110888.00	1108.9	1	8.13E-07	-0.000062923	0.004869	0.090174	-6.977448736	111185.85	-297.85	-0.000242	0.07	8.9936	0.0242					8.967	11.981
			6.18	1.5	-88.826	1460.2	7.30	9.0042	-5.1	111054.41	555.3	1	3.24E-06	-0.000288065	0.025588	0.360151	-31.990917617	111092.10	-37.69	-0.000122	0.00	9.0011	0.0031					8.946	8.555
			6.20	1.6	-101.065	1462.5	0.88	9.0184	-5.8	110879.06	66.5	1	2.26E-04	-0.022832470	2.307558	25.049712	-2531.642896500	110991.89	-112.82	-0.025489	2.88	9.0092	0.0092	<<< same meas. freq. as Lsec.					8.952	9.248
			6.23	1.7	-114.093	1462.4	0.88	9.0177	-6.5	110885.90	66.5	1	2.26E-04	-0.025772194	2.940417	25.047832	-2857.772895734	110885.21	0.69	0.000156	0.00	9.0178	-0.0001					8.942	8.125
			6.26	1.8	-127.910	1462.3	0.88	9.0171	-7.3	110892.69	66.5	1	2.26E-04	-0.028889442	3.695251	25.045940	-3203.628077749	110772.07	120.63	0.027245	3.29	9.0270	-0.0098					8.933	7.188
			6.28	1.9	-142.517	1464.7	0.88	9.0319	-8.1	110710.16	66.4	1	2.27E-04	-0.032294291	4.602488	25.086858	-3575.305996738	110652.46	57.70	0.013074	0.75	9.0366	-0.0047					8.938	7.727
			6.30	2.0	-157.914	1466.8	0.88	9.0449	-9.0	110550.78	66.3	1	2.27E-04	-0.035885829	5.666863	25.122627	-3967.206278407	110526.39	24.39	0.005543	0.14	9.0469	-0.0020					8.940	7.977
			6.34	2.2	-191.076	1471.1	0.88	9.0714	-10.8	110225.76	66.1	1	2.29E-04	-0.043676813	8.345571	25.195845	-4814.309753214	110254.84	-29.09	-0.006648	0.19	9.0690	0.0024					8.944	8.284
			6.38	2.4	-227.396	1474.4	0.88	9.0917	-12.9	109976.99	66.0	1	2.30E-04	-0.052212479	11.872892	25.251893	-5742.171424397	109957.44	19.55	0.004490	0.09	9.0934	-0.0016					8.940	8.046
			6.41	2.6	-266.874	1480.8	0.89	9.1312	-15.1	109499.45	65.7	1	2.32E-04	-0.061810277	16.495562	25.360989	-6768.191196685	109634.17	-134.73	-0.031204	4.20	9.1200	0.0112					8.953	8.637
			6.45	2.8	-309.511	1483.7	0.89	9.1491	-17.4	109283.02	65.6	1	2.33E-04	-0.071966344	22.274360	25.410098	-7864.699596914	109285.05	-2.03	-0.000471	0.00	9.1489	0.0002					8.942	8.139
			6.48	3.0	-355.306	1488.7	0.89	9.1799	-19.9	108913.41	65.4	1	2.34E-04	-0.083172177	29.551554	25.495127	-9058.565582665	108910.06	3.35	0.000785	0.00	9.1802	-0.0003					8.941	8.123
			6.54	3.5	-483.611	1502.4	0.90	9.2644	-26.9	107913.13	64.8	1	2.38E-04	-0.115299765	55.760191	25.728050	-12442.358075526	107859.45	53.68	0.012797	0.69	9.2690	-0.0046					8.937	8.022
			6.60	4.0	-631.655	1520.1	0.91	9.3736	-34.7	106648.45	64.0	1	2.44E-04	-0.154164893	97.378976	26.029171	-16441.447588947	106647.20	1.25	0.000305	0.00	9.3737	-0.0001					8.940	8.131
			6.65	4.5	-799.438	1539.4	0.92	9.4926	-43.4	105302.26	63.2	1	2.50E-04	-0.200100962	159.968304	26.357373	-21071.084075465	105273.33	28.93	0.007242	0.21	9.4952	-0.0026					8.937	8.098
			6.70	5.0	-986.960	1562.9	0.94	9.6375	-52.7	103708.90	62.3	1	2.58E-04	-0.254638212	251.317842	26.757151	-26408.249177007	103737.82	-28.92	-0.007461	0.22	9.6348	0.0027					8.941	8.164
			6.78	6.0	-1421.223	1619.3	0.97	9.9853	-73.3	100074.34	60.1	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	100181.90	-107.57	0.000000	0.00	9.9745	0.0107					8.946	8.213
			6.85	7.0	-1934.442	1698.8	1.02	10.4755	-95.1	95365.85	57.3	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	95979.46	-613.60	0.000000	0.00	10.4086	0.0669					8.985	8.456	Ci computed from Rv
			6.90	8.0	-2526.619	1794.8	1.08	11.0675	-117.6	90237.38	54.2	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	91130.48	-893.10	0.000000	0.00	10.9591	0.1083					9.005	8.495
			6.95	9.0	-3197.752	1920.3	1.15	11.8413	-139.1	84310.79	50.7	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	85634.97	-1324.19	0.000000	0.00	11.6585	0.1828					9.039	8.559
			7.00	10.0	-3947.842	2081.0	1.25	12.8323	-158.4	77769.99	46.8	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	79492.94	-1722.94	0.000000	0.00	12.5547	0.2776					9.070	8.585
			7.04	11.0	-4776.889	2291.0	1.37	14.1272	-174.1	70611.14	42.5	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	72704.37	-2093.23	0.000000	0.00	13.7215	0.4058					9.099	8.590
			7.08	12.0	-5684.892	2574.0	1.54	15.8723	-184.5	62818.29	37.8	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	65269.27	-2450.98	0.000000	0.00	15.2780	0.5944					9.128	8.587
			7.11	13.0	-6671.853	2952.0	1.77	18.2032	-188.8	54746.56	33.0	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	57187.64	-2441.08	0.000000	0.00	17.4288	0.7745					9.127	8.526
			7.15	14.0	-7737.770	3527.0	2.12	21.7489	-183.2	45796.10	27.6	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	48459.49	-2663.39	0.000000	0.00	20.5581	1.1908					9.146	8.509
			7.18	15.0	-8882.644	4432.0	2.66	27.3295	-167.4	36423.11	22.0	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	39084.80	-2661.69	0.000000	0.00	25.4763	1.8532					9.147	8.469
			7.20	16.0	-10106.475	6058.0	3.63	37.3561	-139.3	26630.08	16.1	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	29063.58	-2433.50	0.000000	0.00	34.2432	3.1129					9.130	8.415
			7.23	17.0	-11409.263	9740.0	5.84	60.0608	-97.8	16551.97	10.0	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	18395.83	-1843.86	0.000000	0.00	54.0726	5.9882					9.085	8.341
			7.25	17.6	-12228.835	16157.0	9.69	99.6306	-63.2	9973.87	6.0	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	11684.85	-1710.98	0.000000	0.00	85.1204	14.5101					9.077	8.323
Derived parameters		esd +/-	sums >>>>>>>>>>>>					16	3.28E-03	-1.183067138	672.208286	357.388991	-126785.600982301			1.46E-14	12.73
D	8.07E-01
a	111819.45	27.52
Li / uH	8.942988	0.0022		y obs - y calc
b = Ci / pF	8.188401	0.0608
chi^2/(n-2)	0.909
esd of fit	0.95
Rv0	1450.28
				Incremental error analysis
X-over /MHz	18.6					R0	49.83	49.89	49.95	err contrib.
				Li	8.994662	8.942988	9.016323	0.010831
				Ci	6.173432	8.188401	6.158601	-0.007416

				C2	10.1	10.3	10.5
				Li	8.828914	8.942988	9.182070	0.176578
				Ci	6.289328	8.188401	6.047431	-0.120949

				Accuracy
				esd Li	0.176924
				esd Ci	0.135581

				Li' vs Log(f)

## Sheet2
Variation of C1'm (nominal capacitance, measured @ 1.5915MHz) with frequency					To see details of fitting calculation show columns N-P
			Log(f)	f	Dial	Ci'm	esd C1'm	y	esd y	fit	weight	x	wx	wy	wxy	wx^2	y calc	e	we	we^2		<Ci'm>	obs-calc
				/MHz		/pF		1/C1'm			w					obs-calc
			6.15	1.40	13.90	107.35	0.63	0.009315	0.000054	1	3.40E+08	-22.6	-7690869238	3166934	-71643329.12	1.74E+11	0.009309	0.000006	2033.5	0.012		107.42	-0.069
			6.18	1.50	13.90	107.35	0.13	0.009315	0.000011	1	8.50E+09	-11.2	-94966455885	79173343	-884648125	1.06E+12	0.009310	0.000005	43106.1	0.219		107.41	-0.058
			6.20	1.60	13.90	107.35	0.10	0.009315	0.000009	1	1.33E+10	1.1	14139887179	123708348	131718348	1.51E+10	0.009311	0.000004	54441.7	0.223		107.40	-0.047
			6.23	1.70	13.90	107.35	0.10	0.009315	0.000009	1	1.33E+10	14.1	187150343904	123708348	1743375588	2.64E+12	0.009312	0.000003	40697.1	0.125		107.38	-0.035
			6.26	1.80	13.90	107.35	0.08	0.009315	0.000007	1	2.08E+10	27.9	579134816961	193294294	5394857852	1.62E+13	0.009313	0.000002	40811.8	0.080		107.37	-0.023
			6.28	1.90	13.89	107.32	0.08	0.009318	0.000007	1	2.07E+10	42.5	881407311184	193158962	8212556489	3.75E+13	0.009315	0.000003	61806.7	0.184		107.36	-0.034
			6.30	2.00	13.89	107.32	0.08	0.009318	0.000007	1	2.07E+10	57.9	1200588644773	193158962	11186544450	6.95E+13	0.009316	0.000002	36449.8	0.064		107.34	-0.020
			6.34	2.20	13.87	107.27	0.08	0.009322	0.000007	1	2.07E+10	91.1	1884531912030	192888488	17567423407	1.72E+14	0.009318	0.000003	71942.7	0.250		107.31	-0.040
			6.38	2.40	13.87	107.27	0.08	0.009322	0.000007	1	2.07E+10	127.4	2636067064951	192888488	24573161092	3.36E+14	0.009321	0.000001	12238.1	0.007		107.28	-0.007
			6.41	2.60	13.85	107.22	0.08	0.009326	0.000007	1	2.07E+10	166.9	3446504855335	192618266	32143000390	5.75E+14	0.009324	0.000002	37429.7	0.068		107.24	-0.021
			6.45	2.80	13.83	107.17	0.08	0.009331	0.000007	1	2.06E+10	209.5	4319009791743	192348297	40299044467	9.05E+14	0.009328	0.000003	57438.9	0.160		107.21	-0.032
			6.48	3.00	13.83	107.16	0.08	0.009332	0.000007	1	2.06E+10	255.3	5260600184548	192280844	49090406801	1.34E+15	0.009331	0.000000	4927.3	0.001		107.16	-0.003
			6.54	3.50	13.79	107.07	0.08	0.009339	0.000007	1	2.05E+10	383.6	7878489013581	191809116	73580013136	3.02E+15	0.009342	-0.000002	-47444.3	0.110		107.05	0.026
			6.60	4.00	13.74	106.95	0.08	0.009350	0.000007	1	2.04E+10	531.7	10867958764824	191136558	101618645924	5.78E+15	0.009353	-0.000003	-63979.7	0.200		106.91	0.036
			6.65	4.50	13.70	106.84	0.08	0.009360	0.000007	1	2.04E+10	699.4	14237538707770	190532602	133265733494	9.96E+15	0.009367	-0.000007	-134142.9	0.884		106.76	0.075
			6.70	5.00	13.64	106.70	0.08	0.009372	0.000007	1	2.03E+10	887.0	17961701556471	189796162	168341687392	1.59E+16	0.009382	-0.000009	-190291.4	1.788		106.59	0.107
			6.78	6.00	13.40	106.10	0.08	0.009425	0.000007	1	1.98E+10	1321.2	26157721403024	186604811	246546574622	3.46E+16	0.009416	0.000009	182754.4	1.687		106.20	-0.104
			6.85	7.00	13.27	105.77	0.08	0.009454	0.000007	1	1.96E+10	1834.4	35874489649662	184891197	339172262777	6.58E+16	0.009457	-0.000003	-49119.2	0.123		105.74	0.028
			6.90	8.00	13.06	105.24	0.08	0.009502	0.000007	1	1.92E+10	2426.6	46517752333048	182145266	441997114652	1.13E+17	0.009504	-0.000002	-43766.1	0.100		105.22	0.025
			6.95	9.00	12.83	104.67	0.08	0.009554	0.000007	1	1.88E+10	3097.8	58093066691607	179169164	555021605409	1.80E+17	0.009557	-0.000003	-61500.2	0.202		104.63	0.036
			7.00	10.00	12.56	103.99	0.08	0.009616	0.000007	1	1.83E+10	3847.8	70311941729644	175717058	676131432313	2.71E+17	0.009617	-0.000001	-12957.6	0.009		103.98	0.008
			7.04	11.00	12.24	103.19	0.08	0.009691	0.000008	1	1.77E+10	4676.9	82855540190846	171683440	802944312268	3.88E+17	0.009683	0.000008	144478.6	1.178		103.28	-0.087
			7.08	12.00	12.01	102.61	0.08	0.009745	0.000008	1	1.73E+10	5584.9	96749634800314	168822728	942856728577	5.40E+17	0.009755	-0.000010	-165412.3	1.579		102.51	0.100
			7.11	13.00	11.67	101.76	0.08	0.009827	0.000008	1	1.68E+10	6571.9	110112931238614	164652334	1082070861946	7.24E+17	0.009833	-0.000006	-106614.2	0.678		101.70	0.066
			7.15	14.00	11.32	100.88	0.08	0.009912	0.000008	1	1.62E+10	7637.8	123617460953075	160431611	1225339723225	9.44E+17	0.009918	-0.000006	-90812.2	0.510		100.83	0.057
			7.18	15.00	10.94	99.93	0.08	0.010007	0.000008	1	1.56E+10	8782.6	136855930077026	155931467	1369490554347	1.20E+18	0.010009	-0.000002	-32889.0	0.069		99.91	0.021
			7.20	16.00	10.54	98.93	0.08	0.010108	0.000008	1	1.50E+10	10006.5	149763842898414	151286232	1513841880072	1.50E+18	0.010106	0.000002	30720.4	0.063		98.95	-0.020
			7.23	17.00	10.13	97.90	0.08	0.010214	0.000008	1	1.44E+10	11309.3	162339703558188	146621564	1658181787888	1.84E+18	0.010210	0.000005	66497.6	0.308		97.95	-0.044
			7.25	17.60	9.87	97.25	0.08	0.010283	0.000008	1	1.40E+10	12128.8	169515786950890	143713747	1743080272918	2.06E+18	0.010275	0.000008	111154.8	0.884		97.33	-0.075
			sums >>>>>>>					29	5.05E+11		1340017968004480	4807338630	13262866988392	9.89E+18			-2.57E-07	11.767
Derived parameters		esd					Chi^2(/n-2) >>>>>		0.436
Denom	3.20E+30					esd of fit >>>>>>		0.660
a	0.009311	0.000001					Ci'measured vs Log(f) (with expectation value)
C1' /pF	107.398	0.013
b	7.94E-08	2.62E-10
L1 /nH	79.444	0.262

KT	11.63	< Transformer const. = ( N R0 / k Ri) + (1/N) -1
L2 / nH	50.00	20.00
L1(true)	83.74	1.74	< corrected for L2

			C1'm obs-calc vs Log(f)

## Sheet3
Stray capacitance of voltage sampling network				Predicted C1 = C1' + C1s = C2 [( N R0 / k Ri) + (1/N) -1]
		esd	Deriv dC1/dpar	(der*esd)^2
N =	12
R0 =	49.89	0.06	2.59	0.02
C2 /pF =	10.3	0.2	11.63	5.41
Ri =	49.76	0.06	-2.6	0.02
k =	0.959	0.0112	-134.75	2.28
				7.74	<<< sum
KT	11.629	< Transformer const. = ( N R0 / k Ri) + (1/N) -1

C1' =	107.398	1.074
C1 pred =	119.779	2.781
C1s =	12.382	2.981

C1 prediction for bridge evaluation (see sheet 5)
R0 =	50
C1 pred =	120.064

Self capacitance of transformer secondary				Ci' = Ci apparent + ( Csh * Ro / Ri ) + [ L2 / ( N Ro Rv ) ] + ( LH / Ri^2 ) - Cis - [ L0 / ( R0 Ri ) ] + C0 Ro / Ri
parameter		esd +/-
Csh / pF	0.00	0.00	< shield protrusion capacitance. Set to 0 for gen. side earth.
L2 / nH	50.00	20.00	< inductance of upper voltage sampling arm
L2 fit	1		< =1 when L2 is included in fit on sheet 1, 0 otherwise.
RV0	1450.28		< optimal Rv
LH / nH	6.00	2.00	< inductance of sec. load resistor
Cis / pF	0.40	0.10	< strays accross sec winding
L0 / nH	13.00	3.00	< inductance of pri. transmission line
C0 / pF	3.50	0.50	< capacitance of pri. transmission line

Series for Ci'		esd +/-	Variance
Ci apparent	8.19	0.06	0.004	< from fit on sheet 1
Corr for Csh	0.00	0.00	0.000
Corr for L2	0.00	0.02	0.001	< inclusion depends on whether L2 is included in fit on sheet 1.
Corr for LH	2.42	0.81	0.652
Corr for Cis	-0.40	0.10	0.010
Corr for L0	-5.24	1.21	1.460
Corr for C0	3.51	0.50	0.251
Ci' / pF	8.48	1.54	2.38	Only valid if parameters have been filled in correctly!

lw / mm	228.00	< length of secondary winding wire
Ci' pred / pF	7.64	< predicted Ci' if phase velocity = c

## Sheet4
Data reduction for 8-48 reference capacitor with 68pF in parallel. Data for dial readings 0 and 18 are excluded from fit (end effects).
		Dial	C /pF	esd C	fit	w	wy	wx	wx^2	wxy	y calc	obs-calc	we	we^2
		x	y					C calc	e
		0	74.0	0.20	0	0.000	0.00	0.00	0.00	0.00	72.52	1.48	0.00	0.000
		1	75.0	0.20	1	25.000	1875.00	25.00	25.00	1875.00	75.02	-0.02	-0.59	0.014
		2	77.5	0.20	1	25.000	1937.50	50.00	100.00	3875.00	77.53	-0.03	-0.74	0.022
		3	80.0	0.20	1	25.000	2000.00	75.00	225.00	6000.00	80.04	-0.04	-0.88	0.031
		4	82.5	0.20	1	25.000	2062.50	100.00	400.00	8250.00	82.54	-0.04	-1.03	0.042
		5	85.0	0.20	1	25.000	2125.00	125.00	625.00	10625.00	85.05	-0.05	-1.18	0.055
		6	87.7	0.20	1	25.000	2192.50	150.00	900.00	13155.00	87.55	0.15	3.68	0.541
		7	90.0	0.20	1	25.000	2250.00	175.00	1225.00	15750.00	90.06	-0.06	-1.47	0.087
		8	92.5	0.20	1	25.000	2312.50	200.00	1600.00	18500.00	92.56	-0.06	-1.62	0.105
		9	95.0	0.20	1	25.000	2375.00	225.00	2025.00	21375.00	95.07	-0.07	-1.76	0.125
		10	98.0	0.20	1	25.000	2450.00	250.00	2500.00	24500.00	97.58	0.42	10.59	4.484
		11	100.0	0.20	1	25.000	2500.00	275.00	3025.00	27500.00	100.08	-0.08	-2.06	0.170
		12	102.5	0.20	1	25.000	2562.50	300.00	3600.00	30750.00	102.59	-0.09	-2.21	0.195
		13	105.0	0.20	1	25.000	2625.00	325.00	4225.00	34125.00	105.09	-0.09	-2.35	0.221
		14	108.0	0.20	1	25.000	2700.00	350.00	4900.00	37800.00	107.60	0.40	10.00	4.000
		15	110.0	0.20	1	25.000	2750.00	375.00	5625.00	41250.00	110.11	-0.11	-2.65	0.280
		16	112.5	0.20	1	25.000	2812.50	400.00	6400.00	45000.00	112.61	-0.11	-2.79	0.312
		17	115.0	0.20	1	25.000	2875.00	425.00	7225.00	48875.00	115.12	-0.12	-2.94	0.346
		18	116.0	0.20	0	0.000	0.00	0.00	0.00	0.00	117.62	-1.62	0.00	0.000
		sums >>>>>>>>>			17	425.000	40405.00	3825.00	44625.00	389205.00			-1.39E-11	11.029
D	4335000
a	72.517647		Capacitance vs dial reading
esd a	0.007569
b	2.505882
esd b	0.000072
Chi^2/n-2)	0.735294
esd fit	0.857493

## Sheet5
Bridge performance evaluation - Rvset and Cvset are chosen during 'calibration'.
N	12
R0	50
C2 /pF	10.3
Li /uH	8.943
Rv0	1447.09	<<< Guide value (estimated value for zero phase error at 0Hz)
Rvset	1454.00	<<< Selected value
C1s /pF	12.382
L1 /nH	79.444
Ci pred	120.064					| << - - show hidden columns for details of fit - - >> |
Cv set 0	130.364	<<< Guide value (est. value for zero magnitude error at 0Hz)
C1 set	119.700
Cv set	130.000	<<< Selected value
			x					y				| - - - Determining the gradient of Phi vs f: y=phi, x=f
			f	Rv	C1'm	C1'0	Cv	Xcv	Xcvset	U	Rbal	Xbal	|Zbal|	(|Zbal|/R0)-1	Phi				esd phi	fit	weight					Phi calc	obs-calc
		Log(f)	/MHz	/Ohms	/pF	/pF	/pF				Ohms	Ohms	Ohms	Magn. err.	/deg.					w	wx	wy	wxy	wx^2	Y calc	e	we	we^2
		6.15	1.4	1462.4	107.35	107.26	129.94	-874.9	-874.5	3.46E-07	50.093	-0.117	50.093	0.002	-0.134				0.0200	0	0.00	0.0	0.0	0.0	0.0	-0.087	-0.047	0.000	0.000000
		6.18	1.5	1460.2	107.35	107.26	129.94	-816.6	-816.2	3.59E-07	50.068	-0.081	50.069	0.001	-0.093				0.0200	0	0.00	0.0	0.0	0.0	0.0	-0.101	0.009	0.000	0.000000
		6.20	1.6	1462.5	107.35	107.26	129.94	-765.5	-765.2	3.69E-07	50.081	-0.111	50.081	0.002	-0.126				0.0200	0	0.00	0.0	0.0	0.0	0.0	-0.116	-0.011	0.000	0.000000
		6.23	1.7	1462.4	107.35	107.26	129.94	-720.5	-720.2	3.79E-07	50.075	-0.105	50.075	0.002	-0.120				0.0100	0	0.00	0.0	0.0	0.0	0.0	-0.130	0.010	0.000	0.000000
		6.26	1.8	1462.3	107.35	107.26	129.94	-680.5	-680.1	3.87E-07	50.070	-0.100	50.070	0.001	-0.115				0.0075	0	0.00	0.0	0.0	0.0	0.0	-0.145	0.030	0.000	0.000000
		6.28	1.9	1464.7	107.32	107.23	129.91	-644.8	-644.4	3.94E-07	50.087	-0.124	50.087	0.002	-0.141				0.0075	0	0.00	0.0	0.0	0.0	0.0	-0.159	0.018	0.000	0.000000
		6.30	2.0	1466.8	107.32	107.23	129.91	-612.5	-612.1	4.00E-07	50.093	-0.145	50.094	0.002	-0.166				0.0075	0	0.00	0.0	0.0	0.0	0.0	-0.173	0.008	0.000	0.000000
		6.34	2.2	1471.1	107.27	107.18	129.86	-557.1	-556.5	4.09E-07	50.119	-0.177	50.120	0.002	-0.203				0.0075	1	17777.78	39111.1	-3606.9	-7935.1	86044.4	-0.202	-0.001	-10.831	0.006598
		6.38	2.4	1474.4	107.27	107.18	129.86	-510.6	-510.1	4.17E-07	50.122	-0.201	50.122	0.002	-0.230				0.0075	1	17777.78	42666.7	-4081.0	-9794.3	102400.0	-0.231	0.002	27.890	0.043755
		6.41	2.6	1480.8	107.22	107.13	129.81	-471.5	-470.9	4.22E-07	50.150	-0.246	50.150	0.003	-0.281				0.0075	0	0.00	0.0	0.0	0.0	0.0	-0.260	-0.021	0.000	0.000000
		6.45	2.8	1483.7	107.17	107.08	129.76	-438.0	-437.2	4.27E-07	50.165	-0.253	50.166	0.003	-0.289				0.0075	1	17777.78	49777.8	-5131.1	-14367.0	139377.8	-0.289	0.000	3.385	0.000645
		6.48	3.0	1488.7	107.16	107.07	129.75	-408.9	-408.1	4.30E-07	50.173	-0.280	50.173	0.003	-0.320				0.0075	1	17777.78	53333.3	-5691.2	-17073.5	160000.0	-0.318	-0.002	-43.883	0.108322
		6.54	3.5	1502.4	107.07	106.98	129.66	-350.7	-349.8	4.35E-07	50.209	-0.340	50.210	0.004	-0.388				0.0075	1	17777.78	62222.2	-6893.2	-24126.1	217777.8	-0.390	0.002	36.139	0.073465
		6.60	4.0	1520.1	106.95	106.86	129.54	-307.2	-306.1	4.36E-07	50.260	-0.407	50.262	0.005	-0.464				0.0075	1	17777.78	71111.1	-8243.4	-32973.6	284444.4	-0.462	-0.002	-32.079	0.057884
		6.65	4.5	1539.4	106.84	106.75	129.43	-273.3	-272.1	4.35E-07	50.304	-0.467	50.307	0.006	-0.532				0.0075	1	17777.78	80000.0	-9461.4	-42576.1	360000.0	-0.534	0.002	32.001	0.057603
		6.70	5.0	1562.9	106.70	106.61	129.29	-246.2	-244.9	4.32E-07	50.359	-0.533	50.362	0.007	-0.607				0.0075	1	17777.78	88888.9	-10788.0	-53940.0	444444.4	-0.606	-0.001	-12.623	0.008963
		6.78	6.0	1619.3	106.10	106.01	128.69	-206.1	-204.0	4.22E-07	50.592	-0.648	50.596	0.012	-0.734				0.0075	0	0.00	0.0	0.0	0.0	0.0	-0.750	0.016	0.000	0.000000
		6.85	7.0	1698.8	105.77	105.68	128.36	-177.1	-174.9	4.06E-07	50.721	-0.803	50.727	0.015	-0.906				0.0075	0	0.00	0.0	0.0	0.0	0.0	-0.895	-0.012	0.000	0.000000
		6.90	8.0	1794.8	105.24	105.16	127.84	-155.6	-153.0	3.87E-07	50.927	-0.936	50.935	0.019	-1.053				0.0075	0	0.00	0.0	0.0	0.0	0.0	-1.039	-0.014	0.000	0.000000
		6.95	9.0	1920.3	104.67	104.58	127.26	-139.0	-136.0	3.64E-07	51.153	-1.077	51.165	0.023	-1.206				0.0075	0	0.00	0.0	0.0	0.0	0.0	-1.183	-0.023	0.000	0.000000
		7.00	10.0	2081.0	103.99	103.91	126.59	-125.7	-122.4	3.38E-07	51.421	-1.217	51.436	0.029	-1.356				0.0075	0	0.00	0.0	0.0	0.0	0.0	-1.327	-0.028	0.000	0.000000
		7.04	11.0	2291.0	103.19	103.11	125.79	-115.0	-111.3	3.09E-07	51.743	-1.358	51.761	0.035	-1.503				0.0075	0	0.00	0.0	0.0	0.0	0.0	-1.471	-0.031	0.000	0.000000
		7.08	12.0	2574.0	102.61	102.53	125.21	-105.9	-102.0	2.77E-07	51.974	-1.504	51.996	0.040	-1.657				0.0075	0	0.00	0.0	0.0	0.0	0.0	-1.616	-0.041	0.000	0.000000
		7.11	13.0	2952.0	101.76	101.68	124.36	-98.4	-94.2	2.43E-07	52.322	-1.640	52.348	0.047	-1.796				0.0075	0	0.00	0.0	0.0	0.0	0.0	-1.760	-0.036	0.000	0.000000
		7.15	14.0	3527.0	100.88	100.80	123.49	-92.1	-87.4	2.05E-07	52.685	-1.791	52.715	0.054	-1.947				0.0075	0	0.00	0.0	0.0	0.0	0.0	-1.904	-0.042	0.000	0.000000
		7.18	15.0	4432.0	99.93	99.85	122.53	-86.6	-81.6	1.65E-07	53.084	-1.941	53.120	0.062	-2.094				0.0075	0	0.00	0.0	0.0	0.0	0.0	-2.048	-0.045	0.000	0.000000
		7.20	16.0	6058.0	98.93	98.85	121.53	-81.8	-76.5	1.21E-07	53.511	-2.092	53.552	0.071	-2.238				0.0075	0	0.00	0.0	0.0	0.0	0.0	-2.193	-0.046	0.000	0.000000
		7.23	17.0	9740.0	97.90	97.83	120.51	-77.7	-72.0	7.62E-08	53.956	-2.241	54.003	0.080	-2.379				0.0075	0	0.00	0.0	0.0	0.0	0.0	-2.337	-0.042	0.000	0.000000
		7.25	17.6	16157.0	97.25	97.18	119.86	-75.4	-69.6	4.62E-08	54.242	-2.341	54.293	0.086	-2.471				0.0075	0	0.00	0.0	0.0	0.0	0.0	-2.423	-0.048	0.000	0.000000
					Sums >>>>		8	1.42E+05	487111.1	-53896.0	-202785.6	1794488.9		-3.98E-01	-8.38E-12	0.357234
					Parameter		esd					Chi^2/n-2 >>>>>>>>>>>			0.059539
phase error vs. f					D	1.79E+10					esd of an obs of unit weight >>>			0.244
					a	0.115025	0.002440	<<< phase intercept at 0Hz
					b	-0.144228	0.000687	<<< gradient [degrees / MHz]
					tp	0.4006	0.001908	ns <<< time delay
					lw	0.228	m	<< wire length
					tp pred	0.3803	ns	<< time delay for phase velocity = c

					Phase error, obs-calc

Magnitude error vs f