## Sheet1
FT50-61, 12 turns 0.9mm wire. Lsec = 9.25uH @ 1.5915MHz. 500pF variable ref. cap. No Faraday Shield. 75 Ohm load					D. W. Knight. June 2007.				Version	1.00
		esd					Show columns N-Q to see deatails of fitting procedure.					See sheet 2 for C1'm vs frequency
N =	12.00					Li' = N R0 C2 Rv					Ci computed from Rv
R0 =	75.53	0.06					from fit	from fit
C2/pF	10.95	0.20					Ci =	Li =
L2 / nH	50.00	20.00	< estimated
L2 fit	1	< 1 to include L2 in fit, 0 to exclude it.
1/(N R0 C2)	1.01E+08					1.393	8.969
			Log (f)	f	x	Rv	esd Rv	Li'	Corr(L2)	y	esd y	fit	weight	w x	w x^2	w y	w x y	y calc	obs-calc	w e	w e^2	Li' calc	Li'					Rv Calc	obs-calc	Computed	Computed
				/MHz	-(2 pi f)^2	Ohms	Ohms	/ uH					w					e			/ uH	obs-calc					Li /uH	Ci /pF
			6.20	1.59	-100.000	905.0	0.54	8.9818	-6.1	111330	66.80	1	2.24E-04	-0.022409113	2.240913	24.948081	-2494.809910787	111359	-28.6	-0.006417	0.18	8.9795	0.0023					904.8	0.2	8.971	1.618
			6.23	1.70	-114.093	905.5	0.54	8.9868	-7.0	111268	66.76	1	2.24E-04	-0.025595386	2.920245	24.961672	-2847.942744423	111339	-71.3	-0.016005	1.14	8.9810	0.0058					904.9	0.6	8.974	1.957
			6.26	1.80	-127.910	905.7	0.54	8.9887	-7.8	111242	66.75	1	2.24E-04	-0.028707851	3.672023	24.966997	-3193.530356688	111320	-77.5	-0.017396	1.35	8.9825	0.0063					905.1	0.6	8.974	1.938
			6.28	1.90	-142.517	905.5	0.54	8.9868	-8.7	111266	66.76	1	2.24E-04	-0.031972091	4.556569	24.961284	-3557.409446117	111300	-33.5	-0.007513	0.25	8.9841	0.0027					905.2	0.3	8.971	1.567
			6.30	2.00	-157.914	905.9	0.54	8.9907	-9.6	111216	66.74	1	2.25E-04	-0.035457446	5.599215	24.972100	-3943.435911733	111278	-62.1	-0.013947	0.87	8.9857	0.0050					905.4	0.5	8.973	1.725
			6.34	2.20	-191.076	905.6	0.54	8.9878	-11.6	111251	66.76	1	2.24E-04	-0.042875098	8.192383	24.963377	-4769.890681086	111232	18.9	0.004239	0.08	8.9893	-0.0015					905.8	-0.2	8.966	1.233
			6.38	2.40	-227.396	905.0	0.54	8.9818	-13.9	111322	66.80	1	2.24E-04	-0.050957320	11.587475	24.946341	-5672.690324800	111181	141.0	0.031600	4.46	8.9932	-0.0114					906.1	-1.1	8.956	0.712
			6.41	2.60	-266.874	906.0	0.54	8.9917	-16.2	111197	66.73	1	2.25E-04	-0.059936313	15.995450	24.973366	-6664.744708731	111126	70.7	0.015882	1.12	8.9974	-0.0057					906.6	-0.6	8.962	1.067
			6.45	2.80	-309.511	906.8	0.54	8.9997	-18.8	111096	66.67	1	2.25E-04	-0.069634750	21.552707	24.994834	-7736.170976042	111067	29.4	0.006615	0.19	9.0020	-0.0024					907.0	-0.2	8.965	1.237
			6.48	3.00	-355.306	907.5	0.54	9.0066	-21.6	111008	66.62	1	2.25E-04	-0.080061314	28.446246	25.013502	-8887.441140947	111003	4.7	0.001059	0.00	9.0070	-0.0004					907.5	0.0	8.967	1.319
			6.54	3.50	-483.611	908.5	0.55	9.0165	-29.4	110878	66.54	1	2.26E-04	-0.109212636	52.816390	25.039305	-12109.273823468	110825	53.4	0.012060	0.64	9.0209	-0.0043					908.9	-0.4	8.962	1.221
			6.60	4.00	-631.655	910.2	0.55	9.0334	-38.3	110662	66.42	1	2.27E-04	-0.143179414	90.439947	25.084125	-15844.505195440	110618	43.5	0.009863	0.43	9.0370	-0.0036					910.6	-0.4	8.962	1.263
			6.65	4.50	-799.438	912.6	0.55	9.0572	-48.3	110361	66.25	1	2.28E-04	-0.182168336	145.632283	25.147956	-20104.230222039	110385	-24.0	-0.005467	0.13	9.0553	0.0020					912.4	0.2	8.967	1.362
			6.70	5.00	-986.960	915.0	0.55	9.0810	-59.5	110060	66.07	1	2.29E-04	-0.226083638	223.135607	25.211501	-24882.754198362	110124	-63.6	-0.014573	0.93	9.0758	0.0052					914.5	0.5	8.969	1.397
			6.78	6.00	-1421.223	922.5	0.55	9.1555	-85.0	109139	65.53	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	109519	-379.6	0.000000	0.00	9.1238	0.0317					919.3	3.2	8.993	1.600
			6.85	7.00	-1934.442	933.7	0.56	9.2666	-114.3	107800	64.75	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	108804	-1004.4	0.000000	0.00	9.1812	0.0855					925.1	8.6	9.041	1.853	Li computed from Rv vs Log(f)
			6.90	8.00	-2526.619	946.0	0.57	9.3887	-147.3	106364	63.91	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	107979	-1615.8	0.000000	0.00	9.2484	0.1403					931.9	14.1	9.088	1.974
			6.95	9.00	-3197.752	959.6	0.58	9.5237	-183.8	104818	63.00	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	107045	-2227.2	0.000000	0.00	9.3259	0.1978					939.7	19.9	9.136	2.032
			7.00	10.00	-3947.842	977.4	0.59	9.7003	-222.8	102866	61.85	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	106000	-3133.8	0.000000	0.00	9.4142	0.2862					948.6	28.8	9.209	2.130
			7.08	12.00	-5684.892	1020.0	0.61	10.1231	-307.5	98476	59.27	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	103581	-5104.9	0.000000	0.00	9.6257	0.4974					969.9	50.1	9.372	2.237
			7.15	14.00	-7737.770	1077.0	0.65	10.6888	-396.3	93159	56.13	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	100722	-7563.0	0.000000	0.00	9.8894	0.7995					996.4	80.6	9.585	2.319
			7.20	16.00	-10106.475	1153.0	0.69	11.4431	-483.5	86905	52.43	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	97423	-10518.1	0.000000	0.00	10.2138	1.2293					1029.1	123.9	9.856	2.386
			7.26	18.00	-12791.007	1251.0	0.75	12.4157	-564.0	79979	48.33	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	93685	-13705.9	0.000000	0.00	10.6102	1.8055					1069.1	181.9	10.167	2.420
			7.30	20.00	-15791.367	1393.0	0.84	13.8250	-625.4	71707	43.40	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	89506	-17799.2	0.000000	0.00	11.0949	2.7302					1117.9	275.1	10.602	2.480
			7.34	22.00	-19107.554	1592.0	0.96	15.8000	-662.1	62629	37.97	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	84888	-22259.3	0.000000	0.00	11.6890	4.1110					1177.8	414.2	11.123	2.523
			7.38	24.00	-22739.569	1911.0	1.15	18.9660	-656.4	52070	31.64	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	79830	-27760.6	0.000000	0.00	12.4244	6.5416					1251.9	659.1	11.849	2.585
			7.41	26.00	-26687.410	2478.0	1.49	24.5933	-594.1	40067	24.40	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	74332	-34264.8	0.000000	0.00	13.3464	11.2468					1344.8	1133.2	12.849	2.654
			7.45	28.00	-30951.079	3801.0	2.28	37.7236	-449.2	26059	15.91	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	68394	-42335.0	0.000000	0.00	14.5257	23.1979					1463.6	2337.4	14.365	2.746
			7.48	30.00	-35530.576	8282.0	4.97	82.1959	-236.7	11929	7.30	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	62017	-50087.5	0.000000	0.00	16.0633	66.1325					1618.5	6663.5	16.221	2.796
Derived parameters		esd +/-	sums >>>>>>>>>>>>					14	3.15E-03	-1.108250706	616.787452	350.184440	-122708.829640665			-7.19E-14	11.78
D	7.17E-01
a	111498.05	29.05		y obs - y calc
Li / uH	8.968767	0.0023
b = Ci / pF	1.392636	0.0657
chi^2/(n-2)	0.981850
esd of fit	0.99
Rv0	903.69
				Incremental error analysis
X-over /MHz	45.03					R0	49.83	49.89	49.95	err contrib
				Li	8.959693	8.968767	8.981270	0.010789
				Ci	6.297673	1.392636	6.282543	-0.007565

				C2	10.1	10.3	10.5
				Li	8.796297	8.968767	9.144665	0.174184
				Ci	6.414655	1.392636	6.170287	-0.122184

				Accuracy
				esd Li	0.174533
				esd Ci	0.138936

				Li' vs Log(f)

## Sheet2
Variation of C1'm (nominal capacitance, measured @ 1.5915MHz) with frequency					To see details of fitting calculation show columns N-P
			Log(f)	f	Dial	Ci'm	esd C1'm	y	esd y	fit	weight	x	wx	wy	wxy	wx^2	y calc	e	we	we^2		<Ci'm>	obs-calc
				/MHz		/pF		1/C1'm			w					obs-calc
see sheet 4 for capcal			6.20	1.59	9.885	183.7	0.08	0.005444	0.000002	1	1.78E+11	0	12719807.7	968338875.6	69248.67	9.10E+02	0.005446	-0.000002	-291603.1	0.478		183.63	0.055
			6.23	1.70	9.885	183.7	0.08	0.005444	0.000002	1	1.78E+11	14.09	2506619622166.8	968338875.6	13646438463.94	3.53E+13	0.005447	-0.000003	-473028.5	1.258		183.59	0.090
			6.26	1.80	9.885	183.7	0.08	0.005444	0.000002	1	1.78E+11	27.91	4964293551946.2	968338875.6	27026408743.67	1.39E+14	0.005448	-0.000004	-650912.2	2.382		183.56	0.123
			6.28	1.90	9.885	183.7	0.08	0.005444	0.000002	1	1.78E+11	42.52	7562405991998.7	968338875.6	41170948753.66	3.22E+14	0.005449	-0.000005	-838960.7	3.957		183.52	0.159
			6.30	2.00	9.885	183.7	0.08	0.005444	0.000002	1	1.78E+11	57.91	10300956942324.3	968338875.6	56080058493.93	5.97E+14	0.005450	-0.000006	-1037174.0	6.048		183.49	0.197
			6.34	2.20	9.880	183.6	0.08	0.005448	0.000002	1	1.77E+11	91.08	16158676837307.4	966513740.51	88025761998.83	1.47E+15	0.005452	-0.000005	-852803.0	4.099		183.41	0.162
			6.38	2.40	9.875	183.5	0.08	0.005451	0.000002	1	1.77E+11	127.4	22545802170200.0	964691019.67	122897473650.6	2.87E+15	0.005455	-0.000004	-709082.8	2.841		183.32	0.135
			6.41	2.60	9.875	183.5	0.08	0.005451	0.000002	1	1.77E+11	166.87	29532479862604.1	964691019.67	160981948584.14	4.93E+15	0.005458	-0.000007	-1214770.7	8.338		183.22	0.231
			6.45	2.80	9.865	183.2	0.08	0.005458	0.000002	1	1.76E+11	209.51	36891762177678.9	961052813.93	201350938139.08	7.73E+15	0.005461	-0.000003	-542380.9	1.671		183.12	0.103
			6.48	3.00	9.860	183.1	0.08	0.005461	0.000002	1	1.76E+11	255.31	44842384124187.5	959237325.62	244898812940.86	1.14E+16	0.005464	-0.000003	-518795.5	1.532		183.01	0.099
			6.54	3.50	9.845	182.8	0.08	0.005472	0.000002	1	1.74E+11	383.61	66869834618915.0	953805298.52	365889837778.04	2.57E+16	0.005474	-0.000002	-329856.0	0.624		182.70	0.063
			6.60	4.00	9.830	182.4	0.08	0.005482	0.000002	1	1.73E+11	531.65	91976153342542.8	948394892.36	504218584592.22	4.89E+16	0.005484	-0.000002	-384462.1	0.854		182.34	0.074
			6.65	4.50	9.810	182.0	0.08	0.005496	0.000002	1	1.71E+11	699.44	119782684888554.0	941214570.52	658321195824.19	8.38E+16	0.005496	0.000000	-78968.8	0.036		181.94	0.015
			6.70	5.00	9.790	181.5	0.08	0.005510	0.000002	1	1.70E+11	886.96	150362089846735.0	934072495.06	828485351312	1.33E+17	0.005510	0.000000	-10451.1	0.001		181.49	0.002
			6.78	6.00	9.740	180.3	0.08	0.005545	0.000002	1	1.65E+11	1321.22	218343067101074.0	916383918.37	1210747540715.51	2.88E+17	0.005541	0.000004	616944.2	2.303		180.46	-0.121
			6.85	7.00	9.685	179.1	0.08	0.005584	0.000002	1	1.61E+11	1834.44	294724527174407.0	897199752.8	1645861323984.46	5.41E+17	0.005579	0.000006	936862.2	5.463		179.26	-0.187
			6.90	8.00	9.630	177.8	0.08	0.005624	0.000003	1	1.56E+11	2426.62	378952835270676.0	878299667.73	2131298421353.53	9.20E+17	0.005621	0.000003	428379.8	1.175		177.89	-0.087
			6.95	9.00	9.560	176.2	0.08	0.005676	0.000003	1	1.51E+11	3097.75	466472337918037.0	854652291.64	2647500696981.63	1.45E+18	0.005670	0.000006	838432.9	4.668		176.37	-0.173
			7.00	10.00	9.480	174.4	0.08	0.005735	0.000003	1	1.44E+11	3847.84	555618397770998.0	828179673.31	3186704332103.32	2.14E+18	0.005724	0.000011	1605270.4	17.846		174.69	-0.339
			7.08	12.00	9.310	170.5	0.08	0.005867	0.000003	1	1.32E+11	5584.89	736691921380911.0	773852775.84	4321884281460.39	4.11E+18	0.005850	0.000017	2187344.9	36.271		170.94	-0.483
			7.15	14.00	9.130	166.3	0.08	0.006012	0.000003	1	1.20E+11	7637.77	913621275897719.0	719125574.97	5492515635203.31	6.98E+18	0.005999	0.000013	1578391.0	20.827		166.71	-0.366
			7.20	16.00	8.940	162.0	0.08	0.006173	0.000003	1	1.08E+11	10006.47	1077058265694780.0	664390741.87	6648209286748.57	1.08E+19	0.006170	0.000003	269665.6	0.676		162.07	-0.066
			7.26	18.00	8.720	157.0	0.08	0.006369	0.000003	1	9.50E+10	12691.01	1205098661159940.0	604783437.22	7675311018943.34	1.53E+19	0.006364	0.000005	443523.4	2.072		157.12	-0.115
			7.30	20.00	8.485	151.7	0.08	0.006592	0.000003	1	8.27E+10	15691.37	1298272671827820.0	545423514.1	8558440553009.04	2.04E+19	0.006582	0.000011	881307.5	9.387		151.94	-0.246
			7.34	22.00	8.255	146.5	0.08	0.006825	0.000004	1	7.20E+10	19007.55	1368730196765410.0	491471927.42	9341679259196.51	2.60E+19	0.006822	0.000004	253822.5	0.895		146.59	-0.076
			7.38	24.00	8.025	141.4	0.08	0.007074	0.000004	1	6.24E+10	22639.57	1412888852994680.0	441453430.62	9994315199711.8	3.20E+19	0.007084	-0.000011	-671020.9	7.215		141.15	0.215
			7.41	26.00	7.770	135.7	0.08	0.007370	0.000004	1	5.30E+10	26587.41	1408447440202700.0	390394294.4	10379573284129.4	3.74E+19	0.007370	-0.000001	-34564.7	0.023		135.68	0.012
			7.45	28.00	7.535	130.5	0.08	0.007663	0.000005	1	4.53E+10	30851.08	1398011418658440.0	347244417.79	10712865105006.9	4.31E+19	0.007679	-0.000016	-717567.4	11.363		130.23	0.269
			7.48	30.00	7.290	125.1	0.08	0.007992	0.000005	1	3.83E+10	35430.58	1356734622847640.0	306050102.55	10843531370586.9	4.81E+19	0.008010	-0.000018	-683541.9	12.202		124.84	0.279
			sums >>>>>>>					29	4.05E+12		14693962649362200.0	23094273074.48	98103431137658.5	2.50E+20			4.42E-06	166.506
Derived parameters		esd					Chi^2(/n-2) >>>>>		6.167
Denom	7.95E+32					esd of fit >>>>>>		2.483
a	0.005446	0.000001
C1' /pF	183.628	0.047					Ci'measured vs Log(f) (with expectation value)
b	7.24E-08	1.77E-10
L1 /nH	72.379	0.177

KT	17.67	< Xformer const. = ( N R0 / k Ri) + (1/N) -1
L2 /nH	50	20
L1(true)	75.209	1.146	nH	< corrected for L2

				C1'm obs-calc vs Log(f)

## Sheet3
Stray capacitance of voltage sampling network				Predicted C1 = C1' + C1s = C2 [( N R0 / k Ri) + (1/N) -1]
Parameter		esd +/-	Deriv dC1/dpar	(der*esd)^2
N =	12
R0 =	75.53	0.06	2.69	0.03
C2 /pF =	10.95	0.2	17.67	12.49
Ri =	49.76	0.06	-4.09	0.06
k =	0.980	0.01	-207.67	4.31
				16.89	<<< sum, used to calculate esd of predicted C1.
KT	17.670	< transformer const = ( N R0 / k Ri) + (1/N) -1

C1' =	183.628	1.837
C1 pred =	193.483	4.109
C1s =	9.855	4.501	< NO FARADAY SHIELD

C1 Prediction for bridge evaluation (see sheet 5)
R0 =	75.53
C1 pred =	193.48

Self capacitance of transformer secondary				Ci' = Ci apparent + ( Csh * Ro / Ri ) + [ L2 / ( N Ro Rv ) ] + ( LH / Ri^2 ) - Cis - [ L0 / ( R0 Ri ) ] + C0 Ro / Ri
parameter		esd +/-
Csh / pF	0.00	0.00	< shield protrusion capacitance. Set to 0 for gen. side earth.
L2 / nH	50.00	20.00	< inductance of upper voltage sampling arm
L2 fit	1		< fitting flag =1 when L2 is included in fit on sheet 1, 0 otherwise.
RV0	903.69		< optimal Rv
LH / nH	15.00	2.00	< inductance of sec. load resistor
Cis / pF	0.40	0.10	< strays accross sec winding
L0 / nH	16.00	3.00	< inductance of pri. transmission line
C0 / pF	0.10	0.50	< capacitance of pri. transmission line

Series for Ci'		esd +/-	Variance	NO FARADAY SHIELD
Ci apparent	1.39	0.07	0.004	< from fit on sheet 1
Corr for Csh	0.00	0.00	0.000
Corr for L2	0.00	0.02	0.001	< inclusion depends on whether L2 is included in fit on sheet 1.
Corr for LH	6.06	0.81	0.652
Corr for Cis	-0.40	0.10	0.010
Corr for L0	-4.26	0.80	0.637
Corr for C0	0.15	0.76	0.576
Ci' / pF	2.95	1.37	1.88	Only valid if parameters have been filled in correctly!

lw / mm	228.00	< length of secondary winding wire
Ci' pred / pF	7.64	< predicted Ci' if phase velocity = c

## Sheet4
Data reduction for 500pF ref. cap.					C =	69.83	*exp( -	0.32	* x) -	56.91325	+	24.040788	* x
Non linearity at low settings is corrected by subtracting an exponential function f(x) =( C0-a)Exp(-px)					where x is the dial reading
seed for C0	12.92	<<< seed values are required because the fitting process is iterative
seed for -a	56.91	<<<
C0 - a	69.83					Show columns J-M for details of fitting calculation.
p	0.32
		Dial	f(x)	C meas	esd C	y	fit	w	wy	wx	wx^2	wxy	y calc	obs-calc	we	we^2		C calc	obs-calc	w ec	w (ec)^2
		x		/ pF		C - f(x)					e					ec
0 - 25pF range		0	69.83	15.05	1.06	-54.8	1	0.890	-48.74	0.00	0.00	0.00	-56.91	2.13	1.90	4.049		12.92	2.13	1.898	4.049
		1	50.71	17.40	0.87	-33.3	1	1.322	-44.04	1.32	1.32	-44.04	-32.87	-0.43	-0.57	0.250		17.83	-0.43	-0.575	0.250
0 - 250pF range		2	36.82	28.50	0.91	-8.3	1	1.197	-9.96	2.39	4.79	-19.92	-8.83	0.51	0.61	0.312		27.99	0.51	0.611	0.312
		3	26.74	41.50	0.97	14.8	1	1.072	15.82	3.21	9.64	47.46	15.21	-0.45	-0.48	0.214		41.95	-0.45	-0.479	0.214
		4	19.42	58.00	1.03	38.6	1	0.939	36.23	3.76	15.02	144.92	39.25	-0.67	-0.62	0.416		58.67	-0.67	-0.625	0.416
		5	14.10	76.00	1.10	61.9	1	0.820	50.79	4.10	20.51	253.94	63.29	-1.39	-1.14	1.583		77.39	-1.39	-1.140	1.583
		6	10.24	96.00	1.18	85.8	1	0.713	61.18	4.28	25.68	367.07	87.33	-1.57	-1.12	1.756		97.57	-1.57	-1.119	1.756
		7	7.43	118.00	1.27	110.6	1	0.618	68.34	4.33	30.28	478.35	111.37	-0.81	-0.50	0.402		118.81	-0.81	-0.498	0.402
		8	5.40	141.00	1.36	135.6	1	0.537	72.88	4.30	34.40	583.08	135.41	0.19	0.10	0.019		140.81	0.19	0.101	0.019
		9	3.92	163.00	1.45	159.1	1	0.474	75.45	4.27	38.42	679.09	159.45	-0.37	-0.18	0.066		163.37	-0.37	-0.177	0.066
		10	2.85	187.00	1.55	184.2	1	0.417	76.85	4.17	41.73	768.49	183.49	0.66	0.27	0.181		186.34	0.66	0.275	0.181
		11	2.07	212.00	1.65	209.9	1	0.368	77.30	4.05	44.55	850.27	207.54	2.40	0.88	2.117		209.60	2.40	0.883	2.117
		12	1.50	236.50	1.75	235.0	1	0.328	77.09	3.94	47.24	925.04	231.58	3.42	1.12	3.843		233.08	3.42	1.123	3.843
0 - 2500pF range		13	1.09	260.00	3.54	258.9	1	0.080	20.66	1.04	13.49	268.59	255.62	3.29	0.26	0.865		256.71	3.29	0.263	0.865
		14	0.79	280.00	3.62	279.2	1	0.076	21.31	1.07	14.96	298.29	279.66	-0.45	-0.03	0.015		280.45	-0.45	-0.034	0.015
		15	0.57	302.00	3.71	301.4	1	0.073	21.92	1.09	16.36	328.85	303.70	-2.27	-0.17	0.376		304.27	-2.27	-0.165	0.376
		16	0.42	327.00	3.81	326.6	1	0.069	22.52	1.10	17.65	360.35	327.74	-1.16	-0.08	0.092		328.16	-1.16	-0.080	0.092
		17	0.30	350.00	3.90	349.7	1	0.066	22.99	1.12	19.00	390.85	351.78	-2.08	-0.14	0.285		352.08	-2.08	-0.137	0.285
		18	0.22	375.00	4.00	374.8	1	0.063	23.42	1.13	20.25	421.63	375.82	-1.04	-0.07	0.068		376.04	-1.04	-0.065	0.068
		19	0.16	400.00	4.10	399.8	1	0.059	23.79	1.13	21.48	451.93	399.86	-0.02	0.00	0.000		400.02	-0.02	-0.001	0.000
		20	0.12	425.00	4.20	424.9	1	0.057	24.09	1.13	22.68	481.73	423.90	0.98	0.06	0.055		424.02	0.98	0.056	0.055
		21	0.08	448.00	4.29	447.9	1	0.054	24.32	1.14	23.94	510.62	447.94	-0.03	0.00	0.000		448.03	-0.03	-0.001	0.000
		22	0.06	470.00	4.38	469.9	1	0.052	24.50	1.15	25.23	538.91	471.98	-2.05	-0.11	0.218		472.05	-2.05	-0.107	0.218
		23	0.04	498.00	4.49	498.0	1	0.050	24.68	1.14	26.22	567.60	496.02	1.93	0.10	0.185		496.07	1.93	0.096	0.185
		24	0.03	520.00	4.58	520.0	1	0.048	24.79	1.14	27.46	594.92	520.07	-0.10	0.00	0.000		520.10	-0.10	-0.005	0.000
		25	0.02	542.00	4.67	542.0	1	0.046	24.87	1.15	28.68	621.81	544.11	-2.13	-0.10	0.208		544.13	-2.13	-0.098	0.208
		sums >>>>>>>>>					26	10.489	813.03	58.65	590.98	10869.79			0.00	17.576				0.000	17.576
D	2758.883
a	-56.913247
esd a	0.156871
b	24.040788
esd b	0.002784
Chi^2/n-2)	0.732319
esd fit	0.855756

## Sheet5
Bridge performance evaluation - change Rvset and Cvset to see the effect of different calibration choices.
N	12
R0	75.53
C2 /pF	10.95
Li /uH	8.969
Rv0	903.69	<<< Guide value (estimated value for zero phase error at 0Hz).
Rvset	905.00	<<< selected value (usually set to balance bridge at some chosen 'calibration' frequency, but can be arbitrary)
C1s /pF	9.855
L1 /nH	72.379
C1 pred	193.483					Show columns W-Z for details of fit.
Cv set 0	204.433	<<< Guide value (estimated value for zero magnitude error to zero at 0Hz)
C1 set	193.250
Cv set	204.200	<<< selected value (see note for Rvset)
			x					y				Determining the gradient of Phi vs. f.
			f	Rv	C1'm	C1'0	Cv	Xcv	Xcvset	U	Rbal	Xbal	|Zbal|	(|Zbal|/R0)-1	Phi				esd Phi	fit	weight					Phi calc	obs-calc
		Log(f)	/MHz	Ohms	/pF	/pF	/pF					Magn. err.	deg.				esd y		w	wx	wy	wxy	wx^2	y calc	e	we	we^2
		6.20	1.6	905.0	183.68	183.44	204.24	-489.61	-489.72	9.44E-07	75.52	-0.01	75.52	0.000	-0.005				0.0500	1	400	637	-2	-3	1013	0.107	-0.112	-44.8	5.009
		6.23	1.7	905.5	183.68	183.44	204.24	-458.37	-458.47	9.71E-07	75.53	-0.02	75.53	0.000	-0.018				0.0200	1	2500	4250	-44	-76	7225	0.102	-0.120	-299.6	35.893
		6.26	1.8	905.7	183.68	183.44	204.24	-432.91	-433.00	9.93E-07	75.53	-0.03	75.53	0.000	-0.022				0.0075	1	17778	32000	-393	-707	57600	0.098	-0.120	-2130.7	255.378
		6.28	1.9	905.5	183.68	183.44	204.24	-410.12	-410.21	1.01E-06	75.52	-0.02	75.52	0.000	-0.017				0.0075	1	17778	33778	-295	-560	64178	0.093	-0.110	-1956.7	215.354
		6.30	2.0	905.9	183.68	183.44	204.24	-389.62	-389.70	1.03E-06	75.53	-0.03	75.53	0.000	-0.025				0.0075	1	17778	35556	-448	-897	71111	0.089	-0.114	-2033.8	232.679
		6.34	2.2	905.6	183.57	183.32	204.13	-354.40	-354.28	1.06E-06	75.56	-0.01	75.56	0.000	-0.006				0.0075	1	17778	39111	-110	-241	86044	0.081	-0.087	-1542.9	133.902
		6.38	2.4	905.0	183.45	183.21	204.01	-325.05	-324.75	1.08E-06	75.59	0.02	75.59	0.001	0.017				0.0075	1	17778	42667	295	708	102400	0.072	-0.055	-986.3	54.718
		6.41	2.6	906.0	183.45	183.21	204.01	-300.04	-299.77	1.10E-06	75.60	0.00	75.60	0.001	-0.003				0.0075	1	17778	46222	-59	-153	120178	0.064	-0.067	-1188.0	79.383
		6.45	2.8	906.8	183.22	182.98	203.78	-278.93	-278.36	1.12E-06	75.68	0.00	75.68	0.002	0.001				0.0075	1	17778	49778	15	42	139378	0.055	-0.054	-961.6	52.018
		6.48	3.0	907.5	183.11	182.86	203.67	-260.48	-259.80	1.13E-06	75.73	0.00	75.73	0.003	-0.002				0.0075	1	17778	53333	-42	-125	160000	0.046	-0.049	-866.1	42.192
		6.54	3.5	908.5	182.76	182.52	203.32	-223.65	-222.69	1.15E-06	75.85	0.01	75.85	0.004	0.006				0.0075	1	17778	62222	105	366	217778	0.025	-0.019	-339.5	6.484
		6.60	4.0	910.2	182.41	182.17	202.98	-196.02	-194.85	1.17E-06	75.98	0.00	75.98	0.006	0.003				0.0075	1	17778	71111	57	228	284444	0.004	0.000	-6.4	0.002
		6.65	4.5	912.6	181.95	181.71	202.52	-174.64	-173.20	1.18E-06	76.16	0.00	76.16	0.008	-0.001				0.0075	1	17778	80000	-17	-77	360000	-0.018	0.017	299.9	5.061
		6.70	5.0	915.0	181.49	181.25	202.06	-157.53	-155.88	1.19E-06	76.33	-0.01	76.33	0.011	-0.004				0.0075	1	17778	88889	-75	-376	444444	-0.039	0.035	622.2	21.774
		6.78	6.0	922.5	180.34	180.10	200.91	-132.03	-129.90	1.19E-06	76.77	-0.03	76.77	0.016	-0.023				0.0075	1	17778	106667	-415	-2487	640000	-0.082	0.059	1043.9	61.298
		6.85	7.0	933.7	179.07	178.84	199.64	-113.89	-111.34	1.19E-06	77.26	-0.08	77.26	0.023	-0.060				0.0075	1	17778	124444	-1064	-7447	871111	-0.125	0.065	1155.6	75.114
		6.90	8.0	946.0	177.80	177.58	198.38	-100.28	-97.43	1.19E-06	77.76	-0.13	77.76	0.030	-0.093				0.0075	1	17778	142222	-1656	-13250	1137778	-0.168	0.074	1324.2	98.635
		6.95	9.0	959.6	176.19	175.97	196.77	-89.87	-86.60	1.18E-06	78.40	-0.16	78.40	0.038	-0.116				0.0075	1	17778	160000	-2059	-18527	1440000	-0.210	0.095	1682.9	159.314
		7.00	10.0	977.4	174.36	174.14	194.94	-81.64	-77.94	1.18E-06	79.13	-0.20	79.13	0.048	-0.147				0.0075	1	17778	177778	-2621	-26212	1777778	-0.253	0.106	1881.4	199.098
		7.08	12.0	1020.0	170.46	170.25	191.05	-69.42	-64.95	1.15E-06	80.75	-0.30	80.75	0.069	-0.211				0.0075	1	17778	213333	-3760	-45117	2560000	-0.339	0.127	2264.8	288.516
		7.15	14.0	1077.0	166.34	166.14	186.94	-60.81	-55.67	1.12E-06	82.53	-0.42	82.53	0.093	-0.289				0.0075	1	17778	248889	-5129	-71808	3484444	-0.424	0.136	2417.4	328.726
		7.20	16.0	1153.0	162.01	161.82	182.62	-54.47	-48.71	1.07E-06	84.48	-0.55	84.48	0.119	-0.376				0.0075	1	17778	284444	-6691	-107059	4551111	-0.510	0.134	2377.4	317.930
		7.26	18.0	1251.0	157.01	156.83	177.64	-49.78	-43.30	1.01E-06	86.85	-0.70	86.86	0.150	-0.461				0.0075	1	17778	320000	-8191	-147443	5760000	-0.596	0.135	2399.3	323.809
		7.30	20.0	1393.0	151.70	151.53	172.33	-46.18	-38.97	9.39E-07	89.53	-0.89	89.53	0.185	-0.567				0.0075	1	17778	355556	-10082	-201638	7111111	-0.681	0.114	2030.7	231.962
		7.34	22.0	1592.0	146.52	146.36	167.17	-43.28	-35.43	8.47E-07	92.29	-1.10	92.30	0.222	-0.685				0.0075	1	17778	391111	-12172	-267790	8604444	-0.767	0.082	1462.4	120.289
		7.38	24.0	1911.0	141.37	141.22	162.03	-40.93	-32.48	7.28E-07	95.22	-1.38	95.23	0.261	-0.828				0.0075	1	17778	426667	-14724	-353382	10240000	-0.853	0.024	432.4	10.518
		7.41	26.0	2478.0	135.69	135.56	156.37	-39.15	-29.98	5.82E-07	98.66	-1.71	98.68	0.306	-0.992				0.0075	1	17778	462222	-17637	-458565	12017778	-0.938	-0.054	-958.4	51.669
		7.45	28.0	3801.0	130.50	130.38	151.18	-37.60	-27.84	3.93E-07	102.04	-2.13	102.06	0.351	-1.195				0.0075	1	17778	497778	-21245	-594853	13937778	-1.024	-0.171	-3044.0	521.220
		7.48	30.0	8282.0	125.12	125.01	145.81	-36.38	-25.98	1.87E-07	105.79	-2.57	105.82	0.401	-1.393				0.0075	1	17778	533333	-24758	-742751	16000000	-1.109	-0.283	-5035.7	1426.379
					Sums >>>>		29	482900	5083997.73	-133218.03	-3060200.69	92249127.1			0	5354.32
	Phase error vs f					Parameter		+/- esd					Chi^2/(n-2) >>>>		198.308254
					D	1.87E+13					esd of an obs of unit w >>>>			14.082
					a	0.174802	0.031277	phase intercept at 0Hz
					b	-0.042807	0.002263	degrees / MHz
					tp	0.118908	0.006286	ns (path difference delay)

					Phase err. Deviation from linearity vs. Log(f)

		Magnitude error vs f