## Sheet1
FT50-61, 12 turns 0.9mm wire. Gen side shield earth. 1-turn shift comp winding. Lsec = 9.25uH @ 1.5915MHz.. L-bal coil.					D. W. Knight. June 2007.			Version	1.01
		esd +/-					Show columns N-Q to see deatails of fitting procedure.					See sheet 2 for C1'm vs frequency
N =	12.00					Li' = N R0 C2 Rv
R0 =	49.89	0.06					from fit	from fit
C2+Cx/pF	10.30	0.20	< Adjusted to reproduce Li as measured elsewhere.					Ci =	Li =
L2 / nH	900.00	20.00					Li computed from Rv
L2 fit	1	< 1 to include L2 in fit, 0 to exclude it.
1/(N R0 C2)	1.62E+08					0.687	9.718
			Log (f)	f	x	Rv	esd Rv	Li'	Corr(L2)	y	esd y	fit	weight	w x	w x^2	w y	w x y	y calc	obs-calc	w e	w e^2	Li' calc	Li'					Computed	Computed
				/MHz	-(2 pi f)^2	Ohms	Ohms	/ uH					w					e			/ uH	obs-calc					Li /uH	Ci /pF
			6.20	1.59	-100.000	1576.5	0.95	9.7213	-95.4	102771.16	61.7	1	2.63E-04	-0.026251233	2.625125	26.978678	-2697.869768536	102835.78	-64.62	-0.016963	1.10	9.7152	0.0061					9.715	0.380
			6.23	1.70	-114.093	1576.0	0.95	9.7183	-108.8	102790.33	61.7	1	2.62E-04	-0.029931705	3.414987	26.966595	-3076.689694528	102826.09	-35.77	-0.009384	0.34	9.7149	0.0034					9.711	0.047
			6.26	1.80	-127.910	1575.4	0.95	9.7146	-122.1	102816.29	61.8	1	2.62E-04	-0.033531106	4.288966	26.952872	-3447.543875861	102816.59	-0.31	-0.000080	0.00	9.7145	0.0000					9.706	-0.264
			6.28	1.90	-142.517	1575.7	0.95	9.7164	-136.0	102782.78	61.8	1	2.62E-04	-0.037374506	5.326506	26.954350	-3841.455505778	102806.55	-23.77	-0.006235	0.15	9.7142	0.0022					9.707	-0.100
Calibration freq. >>>			6.30	2.00	-157.914	1575.3	0.95	9.7139	-150.7	102794.18	61.8	1	2.62E-04	-0.041391172	6.536232	26.943657	-4254.771710452	102795.97	-1.78	-0.000468	0.00	9.7138	0.0002					9.704	-0.256
			6.34	2.20	-191.076	1574.3	0.94	9.7078	-182.5	102827.81	61.8	1	2.62E-04	-0.050019752	9.557551	26.918263	-5143.421699244	102773.17	54.64	0.014304	0.78	9.7129	-0.0052					9.695	-0.553
			6.38	2.40	-227.396	1574.4	0.94	9.7084	-217.1	102786.60	61.8	1	2.62E-04	-0.059535202	13.538048	26.910893	-6119.421025865	102748.20	38.40	0.010054	0.39	9.7120	-0.0036					9.694	-0.436
			6.41	2.60	-266.874	1574.3	0.94	9.7078	-254.8	102755.43	61.8	1	2.62E-04	-0.069862299	18.644438	26.899315	-7178.730667948	102721.06	34.37	0.008998	0.31	9.7110	-0.0032					9.691	-0.396
			6.45	2.80	-309.511	1575.2	0.95	9.7133	-295.4	102656.03	61.8	1	2.62E-04	-0.081116397	25.106401	26.904029	-8327.087398399	102691.75	-35.72	-0.009360	0.33	9.7100	0.0034					9.693	-0.151
			6.48	3.00	-355.306	1574.4	0.94	9.7084	-339.3	102664.47	61.8	1	2.62E-04	-0.093023753	33.051875	26.878917	-9550.233967775	102660.26	4.20	0.001100	0.00	9.7088	-0.0004					9.685	-0.279
			6.54	3.50	-483.611	1572.8	0.94	9.6985	-462.2	102646.27	61.9	1	2.61E-04	-0.126358446	61.108286	26.819559	-12970.223200757	102572.06	74.21	0.019391	1.44	9.7055	-0.0070					9.667	-0.422
			6.60	4.00	-631.655	1573.3	0.94	9.7016	-603.6	102472.19	61.8	1	2.61E-04	-0.165144553	104.314330	26.791101	-16922.724210894	102470.28	1.91	0.000500	0.00	9.7018	-0.0002					9.661	-0.271
			6.65	4.50	-799.438	1572.9	0.94	9.6991	-764.1	102337.89	61.9	1	2.61E-04	-0.208904809	167.006434	26.742385	-21378.877586269	102354.93	-17.04	-0.004453	0.08	9.6975	0.0016					9.648	-0.247
			6.70	5.00	-986.960	1572.3	0.94	9.6954	-943.7	102197.65	61.9	1	2.61E-04	-0.257710446	254.350015	26.685367	-26337.401944666	102226.01	-28.36	-0.007406	0.21	9.6928	0.0027					9.632	-0.240
			6.78	6.00	-1421.223	1573.7	0.94	9.7041	-1357.7	101691.89	61.8	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	101927.46	-235.57	0.000000	0.00	9.6819	0.0221					9.613	-0.102
			6.85	7.00	-1934.442	1588.5	0.95	9.7953	-1830.7	100258.74	61.3	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	101574.63	-1315.89	0.000000	0.00	9.6707	0.1247					9.669	0.421
			6.90	8.00	-2526.619	1587.9	0.95	9.7916	-2392.0	99735.99	61.3	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	101167.52	-1431.53	0.000000	0.00	9.6563	0.1354					9.628	0.307
			6.95	9.00	-3197.752	1590.2	0.95	9.8058	-3023.0	98957.28	61.2	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	100706.13	-1748.85	0.000000	0.00	9.6405	0.1653					9.599	0.289	Ci computed from Rv
			7.00	10.00	-3947.842	1597.4	0.96	9.8502	-3715.3	97805.34	60.9	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	100190.45	-2385.11	0.000000	0.00	9.6241	0.2261					9.594	0.351
			7.08	12.00	-5684.892	1611.4	0.97	9.9365	-5303.5	95335.07	60.4	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	98996.26	-3661.19	0.000000	0.00	9.5877	0.3488					9.565	0.399
			7.15	14.00	-7737.770	1633.1	0.98	10.0704	-7122.8	92178.57	59.6	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	97584.94	-5406.37	0.000000	0.00	9.5504	0.5200					9.558	0.466
			7.20	16.00	-10106.475	1624.2	0.97	10.0155	-9354.2	90491.28	59.9	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	95956.50	-5465.22	0.000000	0.00	9.4957	0.5198					9.364	0.303
			7.26	18.00	-12791.007	1614.7	0.97	9.9569	-11908.6	88524.34	60.3	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	94110.93	-5586.58	0.000000	0.00	9.4322	0.5247					9.155	0.193
			7.30	20.00	-15791.367	1595.7	0.96	9.8397	-14877.0	86751.77	61.0	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	92048.23	-5296.46	0.000000	0.00	9.3523	0.4874					8.890	0.081
			7.34	22.00	-19107.554	1590.2	0.95	9.8058	-18063.5	83916.83	61.2	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	89768.40	-5851.57	0.000000	0.00	9.2737	0.5321					8.687	0.048
Calibration freq. >>>			7.38	24.00	-22739.569	1576.0	0.95	9.7183	-21690.7	81208.45	61.7	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	87271.45	-6063.01	0.000000	0.00	9.1775	0.5408					8.437	0.000
			7.41	26.00	-26687.410	1562.6	0.94	9.6356	-25674.8	78106.80	62.3	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	84557.38	-6450.58	0.000000	0.00	9.0718	0.5639					8.188	-0.033
			7.45	28.00	-30951.079	1555.5	0.93	9.5918	-29912.6	74342.71	62.6	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	81626.18	-7283.47	0.000000	0.00	8.9655	0.6263					7.966	-0.044
			7.48	30.00	-35530.576	1546.9	0.93	9.5388	-34529.3	70305.58	62.9	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	78477.85	-8172.27	0.000000	0.00	8.8490	0.6898					7.736	-0.054
Derived parameters		esd +/-	sums >>>>>>>>>>>>					14	3.67E-03	-1.280155379	708.869194	376.345982	-131246.452256974			5.46E-15	5.12
D	9.60E-01
a	102904.53	17.76
Li / uH	9.717745	0.0017		y obs - y calc
b = Ci / pF	0.687483	0.0404
chi^2/(n-2)	0.427
esd of fit	0.65
Rv0	1575.92
				Incremental error analysis
X-over /MHz	61.58					R0	49.83	49.89	49.95	err contrib.
				Li	8.994662	9.717745	9.016323	0.010831
				Ci	6.173432	0.687483	6.158601	-0.007416

				C2	10.1	10.3	10.5
				Li	8.828914	9.717745	9.182070	0.176578
				Ci	6.289328	0.687483	6.047431	-0.120949

				Accuracy
				esd Li	0.176918
				esd Ci	0.127726

				Li' vs Log(f)

## Sheet2
Variation of C1'm (nominal capacitance, measured @ 1.5915MHz) with frequency					To see details of fitting calculation show columns N-P
			Log(f)	f	Dial	Ci'm	esd C1'm	y	esd y	fit	weight	x	wx	wy	wxy	wx^2	y calc	e	we	we^2		<Ci'm>	obs-calc
				/MHz		/pF		1/C1'm			w					obs-calc
			6.20	1.59	5.430	100.72	0.63	0.009928	0.000062	1	2.63E+08	0.0	18841	2615724	187.06	1.35E+00	0.009932	-0.000004	-989.7	0.004		100.68	0.038
			6.23	1.70	5.430	100.72	0.13	0.009928	0.000012	1	6.59E+09	14.1	92819908696	65393093	921560464	1.31E+12	0.009932	-0.000004	-24805.1	0.093		100.68	0.038
			6.26	1.80	5.430	100.72	0.10	0.009928	0.000010	1	1.03E+10	27.9	287230254093	102176708	2851759389	8.02E+12	0.009932	-0.000004	-38852.4	0.147		100.68	0.038
			6.28	1.90	5.420	100.69	0.10	0.009931	0.000010	1	1.03E+10	42.5	437113567499	102099375	4340968075	1.86E+13	0.009932	-0.000001	-13148.2	0.017		100.68	0.013
			6.30	2.00	5.420	100.69	0.08	0.009931	0.000008	1	1.61E+10	57.9	930319010427	159530274	9238983696	5.39E+13	0.009932	-0.000001	-20708.3	0.027		100.68	0.013
			6.34	2.20	5.410	100.67	0.08	0.009933	0.000008	1	1.60E+10	91.1	1461551170717	159409502	14518306662	1.33E+14	0.009932	0.000001	19196.2	0.023		100.68	-0.012
			6.38	2.40	5.410	100.67	0.08	0.009933	0.000008	1	1.60E+10	127.4	2044405234145	159409502	20308082759	2.60E+14	0.009932	0.000001	18809.3	0.022		100.68	-0.012
			6.41	2.60	5.400	100.64	0.08	0.009936	0.000008	1	1.60E+10	166.9	2675238816364	159288791	26581174130	4.46E+14	0.009932	0.000004	58586.7	0.214		100.68	-0.037
			6.45	2.80	5.400	100.64	0.08	0.009936	0.000008	1	1.60E+10	209.5	3358768068323	159288791	33372721097	7.04E+14	0.009932	0.000004	58132.9	0.211		100.68	-0.037
			6.48	3.00	5.395	100.63	0.08	0.009937	0.000008	1	1.60E+10	255.3	4090862251503	159228459	40651942365	1.04E+15	0.009932	0.000005	77722.2	0.377		100.68	-0.049
			6.54	3.50	5.390	100.62	0.08	0.009939	0.000008	1	1.60E+10	383.6	6143635871475	159168141	61058588628	2.36E+15	0.009932	0.000006	96419.4	0.580		100.68	-0.061
			6.60	4.00	5.380	100.59	0.08	0.009941	0.000008	1	1.60E+10	531.7	8506004731799	159047552	84558375739	4.52E+15	0.009933	0.000008	134925.6	1.138		100.68	-0.085
			6.65	4.50	5.380	100.59	0.08	0.009941	0.000008	1	1.60E+10	699.4	11190388747832	159047552	111243894900	7.83E+15	0.009933	0.000008	133143.4	1.108		100.68	-0.084
			6.70	5.00	5.370	100.57	0.08	0.009944	0.000008	1	1.60E+10	887.0	14176246079326	158927024	140961983295	1.26E+16	0.009933	0.000011	171174.8	1.833		100.68	-0.108
			6.78	6.00	5.350	100.52	0.08	0.009949	0.000008	1	1.60E+10	1321.2	21074381064005	158686151	209659797545	2.78E+16	0.009933	0.000015	246440.3	3.808		100.67	-0.156
			6.85	7.00	5.330	100.47	0.08	0.009954	0.000008	1	1.59E+10	1834.4	29201424754579	158445521	290659191545	5.36E+16	0.009933	0.000020	320647.1	6.459		100.67	-0.204
			6.90	8.00	5.395	100.63	0.08	0.009937	0.000008	1	1.60E+10	2426.6	38882644122276	159228459	386386759238	9.44E+16	0.009934	0.000003	54624.0	0.186		100.67	-0.035
			6.95	9.00	5.385	100.61	0.08	0.009940	0.000008	1	1.60E+10	3097.8	49586339496478	159107839	492876599016	1.54E+17	0.009934	0.000005	87622.7	0.480		100.66	-0.055
			7.00	10.00	5.400	100.64	0.08	0.009936	0.000008	1	1.60E+10	3847.8	61686597568762	159288791	612918062286	2.37E+17	0.009935	0.000001	19409.1	0.023		100.66	-0.012
			7.08	12.00	5.435	100.73	0.08	0.009927	0.000008	1	1.61E+10	5584.9	89851058702755	159711546	891971756330	5.02E+17	0.009936	-0.000009	-140243.0	1.223		100.64	0.088
			7.15	14.00	5.470	100.82	0.08	0.009918	0.000008	1	1.61E+10	7637.8	123312873463233	160135048	1223074641398	9.42E+17	0.009937	-0.000019	-304160.1	5.730		100.63	0.191
			7.20	16.00	5.495	100.89	0.08	0.009912	0.000008	1	1.62E+10	10006.5	161963609925978	160438007	1605418895140	1.62E+18	0.009939	-0.000027	-431495.0	11.503		100.61	0.271
			7.26	18.00	5.485	100.86	0.08	0.009915	0.000008	1	1.62E+10	12691.0	205208203960309	160316778	2034581398211	2.60E+18	0.009941	-0.000026	-419489.2	10.883		100.60	0.263
			7.30	20.00	5.470	100.82	0.08	0.009918	0.000008	1	1.61E+10	15691.4	253339338101234	160135048	2512737813968	3.98E+18	0.009943	-0.000024	-390484.5	9.444		100.58	0.245
			7.34	22.00	5.450	100.77	0.08	0.009923	0.000008	1	1.61E+10	19007.6	306261213940087	159892955	3039173999719	5.82E+18	0.009945	-0.000021	-344555.1	7.368		100.55	0.217
			7.38	24.00	5.405	100.66	0.08	0.009935	0.000008	1	1.60E+10	22639.6	363129137785254	159349139	3607595751688	8.22E+18	0.009947	-0.000013	-200808.8	2.514		100.53	0.127
			7.41	26.00	5.340	100.49	0.08	0.009951	0.000008	1	1.59E+10	26587.4	423658001085266	158565805	4215854127985	1.13E+19	0.009950	0.000001	18996.9	0.023		100.50	-0.012
			7.45	28.00	5.270	100.31	0.08	0.009969	0.000008	1	1.58E+10	30851.1	488125396406759	157725091	4865989307903	1.51E+19	0.009953	0.000016	253322.7	4.056		100.48	-0.161
			7.48	30.00	5.180	100.08	0.08	0.009992	0.000008	1	1.57E+10	35430.6	555486309865959	156648546	5550148193356	1.97E+19	0.009956	0.000036	560566.2	20.043		100.44	-0.359
			sums >>>>>>>					29	4.28E+11		3226161113973970	4252305212	32089654636714	7.03E+19			1.93E-07	89.536
Derived parameters		esd					Chi^2(/n-2) >>>>>		3.316
Denom	1.97E+31					esd of fit >>>>>>		1.821
a	0.009932	0.000003					Ci'measured vs Log(f) (with expectation value)
C1' /pF	100.682	0.035
b	6.64E-10	2.69E-10
L1 /nH	0.664	0.269

KT	10.99	< Transformer const. = ( N R0 / k Ri) + (1/N) -1
L2 / nH	900.00	20.00
L1(true)	82.53	1.84	< corrected for L2

				C1'm obs-calc vs Log(f)

## Sheet3
Stray capacitance of voltage sampling network				Predicted C1 = C1' + C1s = C2 [( (1+Cx/C2) N R0 / k Ri ) + (1/N) -1]
		esd	Deriv dC1/dpar	(der*esd)^2
N =	12
R0 =	49.89	0.06	2.46	0.02
C2+Cx /pF =	10.3	0.2	10.99	4.83
Cx /pF	0	0
C2 /pF	10.3
1+Cx/C2	1
Ri =	49.77	0.06	-2.46	0.02
k =	1.010	0.0200	-121.46	5.9
				10.78	<<< sum
KT	10.993	< Transformer const. = [ (1+Cx/C2) N R0 / k Ri ] + (1/N) -1

C1' =	100.682	1.007
C1 pred =	113.230	3.283
C1s =	12.547	3.434

C1 prediction for bridge evaluation (see sheet 5)
R0 =	50
C1 pred =	113.500

Self capacitance of transformer secondary				Ci' = Ci apparent + ( Csh * Ro / Ri ) + [ L2 / ( N Ro Rv ) ] + ( LH / Ri^2 ) - Cis - [ L0 / ( R0 Ri ) ] + (C0 Ro / Ri) - [(N-1)Cx R0/Ri] -Cx C2/(N(Cx+C2)
parameter		esd +/-
Cqi / pF	0.00	0.00	< neutralising capacitor from load terminal to 1 turn current injection winding // 21.7 Ohms..
L2 / nH	900.00	20.00	< inductance of upper voltage sampling arm
L2 fit	1		< =1 when L2 is included in fit on sheet 1, 0 otherwise.
RV0	1575.92		< optimal Rv
LH / nH	15.00	2.00	< inductance of sec. load resistor
Cis / pF	0.40	0.10	< strays accross sec winding
L0 / nH	13.00	3.00	< inductance of pri. transmission line
C0 / pF	3.50	0.50	< capacitance of pri. transmission line
Cn0 /pF	2.95	0.20	< neutralising capacitor across C0 = Cqi

Series for Ci'		esd +/-	Variance
Ci apparent	0.69	0.04	0.002	< from fit on sheet 1
Corr for Cqi	0.00	0.00	0.000
Corr for Cx	0.00	0.00	0.000
Corr for L2	0.00	0.02	0.000	< inclusion depends on whether L2 is included in fit on sheet 1.
Corr for LH	6.06	0.81	0.652
Corr for Cis	-0.40	0.10	0.010
Corr for L0	-5.24	1.21	1.460
Corr for C0	6.47	0.54	0.291
Ci' / pF	7.57	1.55	2.42	Only valid if parameters have been filled in correctly!

lw / mm	228.00	< length of secondary winding wire
Ci' pred / pF	7.64	< predicted Ci' if phase velocity = c

## Sheet4
Data reduction for 8-48 reference capacitor with 82pF in parallel. Data for dial readings 0 and 18 are excluded from fit (end effects).
		Dial	C /pF	esd C	fit	w	wy	wx	wx^2	wxy	y calc	obs-calc	we	we^2
		x	y					C calc	e
		0	89.0	0.20	0	0.000	0.00	0.00	0.00	0.00	86.92	2.08	0.00	0.000
		1	90.0	0.20	1	25.000	2250.00	25.00	25.00	2250.00	89.46	0.54	13.48	7.269
		2	92.0	0.20	1	25.000	2300.00	50.00	100.00	4600.00	92.00	0.00	-0.06	0.000
		3	94.5	0.20	1	25.000	2362.50	75.00	225.00	7087.50	94.54	-0.04	-1.10	0.049
		4	97.0	0.20	1	25.000	2425.00	100.00	400.00	9700.00	97.09	-0.09	-2.14	0.184
		5	99.5	0.20	1	25.000	2487.50	125.00	625.00	12437.50	99.63	-0.13	-3.19	0.406
		6	102.0	0.20	1	25.000	2550.00	150.00	900.00	15300.00	102.17	-0.17	-4.23	0.715
		7	104.5	0.20	1	25.000	2612.50	175.00	1225.00	18287.50	104.71	-0.21	-5.27	1.111
		8	107.0	0.20	1	25.000	2675.00	200.00	1600.00	21400.00	107.25	-0.25	-6.31	1.593
		9	110.0	0.20	1	25.000	2750.00	225.00	2025.00	24750.00	109.79	0.21	5.15	1.060
		10	112.0	0.20	1	25.000	2800.00	250.00	2500.00	28000.00	112.34	-0.34	-8.39	2.819
		11	115.0	0.20	1	25.000	2875.00	275.00	3025.00	31625.00	114.88	0.12	3.06	0.375
		12	117.5	0.20	1	25.000	2937.50	300.00	3600.00	35250.00	117.42	0.08	2.02	0.164
		13	120.0	0.20	1	25.000	3000.00	325.00	4225.00	39000.00	119.96	0.04	0.98	0.038
		14	122.5	0.20	1	25.000	3062.50	350.00	4900.00	42875.00	122.50	0.00	-0.06	0.000
		15	125.0	0.20	1	25.000	3125.00	375.00	5625.00	46875.00	125.04	-0.04	-1.10	0.049
		16	128.0	0.20	1	25.000	3200.00	400.00	6400.00	51200.00	127.59	0.41	10.36	4.289
		17	130.0	0.20	1	25.000	3250.00	425.00	7225.00	55250.00	130.13	-0.13	-3.19	0.406
		18	131.0	0.20	0	0.000	0.00	0.00	0.00	0.00	132.67	-1.67	0.00	0.000
		sums >>>>>>>>>			17	425.000	46662.50	3825.00	44625.00	445887.50			-1.49E-11	20.527
D	4335000
a	86.919118		Capacitance vs dial reading
esd a	0.014087
b	2.541667
esd b	0.000134
Chi^2/n-2)	1.368464
esd fit	1.169814

## Sheet5
Bridge performance evaluation				FT50-61, 12 turns 0.9mm wire. Gen side shield earth. 1-turn shift comp winding. Lsec = 9.25uH @ 1.5915MHz.. L-bal coil.
N	12
R0	50
C2 /pF	10.3
Li /uH	9.718
Rv0	1572.45	<<< Guide value (estimated value for zero phase error at 0Hz)
Rvset	1578.00	<<< Selected value
C1s /pF	12.547
L1 /nH	0.664
C1 pred	113.500								| << - - show hidden columns for details of fit - - >> |
Cv set 0	123.800	<<< Guide value (est. value for zero magnitude error at 0Hz)
C1 set	113.020
Cv set	123.320	<<< Selected value
			x						y				| - - - Determining the gradient of Phi vs f: y=phi, x=f
			f	Rv	C1'm	C1'0	Cv	Xcv	Xcvset	U	Rbal	Xbal	|Zbal|	(|Zbal|/R0)-1	Phi				esd phi	fit	weight					Phi calc	obs-calc
		Log(f)	/MHz	/Ohms	/pF	/pF	/pF				Ohms	Ohms	Ohms	Magn. err.	/deg.					w	wx	wy	wxy	wx^2	Y calc	e	we	we^2
		6.20	1.6	1576.5	100.72	100.72	123.57	-809.3	-810.9	3.18E-07	49.911	-0.021	49.911	-0.0018	-0.024				0.0500	1	400.00	636.6	-9.8	-15.6	1013.2	0.009	-0.034	-13.432	0.451028
		6.23	1.7	1576.0	100.72	100.72	123.57	-757.6	-759.2	3.26E-07	49.907	-0.014	49.907	-0.0019	-0.016				0.0500	1	400.00	680.0	-6.6	-11.2	1156.0	0.009	-0.025	-10.102	0.255106
		6.26	1.8	1575.4	100.72	100.72	123.57	-715.6	-717.0	3.33E-07	49.903	-0.007	49.903	-0.0019	-0.008				0.0500	1	400.00	720.0	-3.0	-5.5	1296.0	0.009	-0.016	-6.479	0.104930
		6.28	1.9	1575.7	100.69	100.69	123.54	-678.0	-679.3	3.39E-07	49.913	-0.006	49.913	-0.0017	-0.007				0.0500	1	400.00	760.0	-2.8	-5.3	1444.0	0.008	-0.015	-6.141	0.094290
Calibration freq. >>>		6.30	2.0	1575.3	100.69	100.69	123.54	-644.1	-645.3	3.44E-07	49.911	-0.001	49.911	-0.0018	-0.002				0.0500	1	400.00	800.0	-0.7	-1.3	1600.0	0.008	-0.010	-3.903	0.038080
		6.34	2.2	1574.3	100.67	100.67	123.52	-585.7	-586.6	3.53E-07	49.916	0.012	49.916	-0.0017	0.014				0.0500	1	400.00	880.0	5.7	12.5	1936.0	0.008	0.007	2.640	0.017425
		6.38	2.4	1574.4	100.67	100.67	123.52	-536.9	-537.7	3.60E-07	49.917	0.011	49.917	-0.0017	0.012				0.0500	1	400.00	960.0	4.9	11.7	2304.0	0.007	0.005	2.028	0.010283
		6.41	2.6	1574.3	100.64	100.64	123.49	-495.7	-496.4	3.66E-07	49.927	0.014	49.927	-0.0015	0.016				0.0500	1	400.00	1040.0	6.3	16.4	2704.0	0.007	0.009	3.695	0.034134
		6.45	2.8	1575.2	100.64	100.64	123.49	-460.3	-460.9	3.70E-07	49.929	0.005	49.929	-0.0014	0.006				0.0500	1	400.00	1120.0	2.4	6.8	3136.0	0.006	0.000	-0.007	0.000000
		6.48	3.0	1574.4	100.63	100.63	123.48	-429.6	-430.2	3.74E-07	49.933	0.013	49.933	-0.0013	0.015				0.0500	1	400.00	1200.0	5.8	17.5	3600.0	0.006	0.009	3.612	0.032621
		6.54	3.5	1572.8	100.62	100.62	123.47	-368.3	-368.7	3.82E-07	49.936	0.024	49.936	-0.0013	0.027				0.0500	1	400.00	1400.0	10.8	37.8	4900.0	0.004	0.023	9.073	0.205790
		6.60	4.0	1573.3	100.59	100.59	123.44	-322.3	-322.6	3.86E-07	49.947	0.020	49.947	-0.0011	0.023				0.0500	1	400.00	1600.0	9.1	36.2	6400.0	0.003	0.020	7.842	0.153754
		6.65	4.5	1572.9	100.59	100.59	123.44	-286.5	-286.8	3.90E-07	49.948	0.020	49.948	-0.0010	0.023				0.0500	1	400.00	1800.0	9.1	41.1	8100.0	0.002	0.021	8.439	0.178034
		6.70	5.0	1572.3	100.57	100.57	123.41	-257.9	-258.1	3.92E-07	49.958	0.023	49.958	-0.0008	0.026				0.0500	1	400.00	2000.0	10.4	52.1	10000.0	0.000	0.026	10.232	0.261714
		6.78	6.0	1573.7	100.52	100.52	123.36	-215.0	-215.1	3.95E-07	49.980	0.016	49.980	-0.0004	0.018				0.0500	0	0.00	0.0	0.0	0.0	0.0	-0.002	0.020	0.000	0.000000
		6.85	7.0	1588.5	100.47	100.47	123.31	-184.4	-184.4	3.94E-07	50.007	-0.038	50.007	0.0001	-0.043				0.0500	0	0.00	0.0	0.0	0.0	0.0	-0.005	-0.039	0.000	0.000000
		6.90	8.0	1587.9	100.63	100.63	123.48	-161.1	-161.3	3.95E-07	49.940	-0.038	49.940	-0.0012	-0.044				0.0500	0	0.00	0.0	0.0	0.0	0.0	-0.007	-0.036	0.000	0.000000
		6.95	9.0	1590.2	100.61	100.61	123.45	-143.2	-143.4	3.95E-07	49.950	-0.039	49.950	-0.0010	-0.045				0.0500	0	0.00	0.0	0.0	0.0	0.0	-0.010	-0.035	0.000	0.000000
		7.00	10.0	1597.4	100.64	100.64	123.49	-128.9	-129.1	3.94E-07	49.935	-0.055	49.935	-0.0013	-0.063				0.0500	0	0.00	0.0	0.0	0.0	0.0	-0.012	-0.051	0.000	0.000000
		7.08	12.0	1611.4	100.73	100.73	123.58	-107.3	-107.5	3.91E-07	49.900	-0.077	49.900	-0.0020	-0.089				0.0500	0	0.00	0.0	0.0	0.0	0.0	-0.017	-0.071	0.000	0.000000
		7.15	14.0	1633.1	100.82	100.82	123.67	-91.9	-92.2	3.86E-07	49.865	-0.106	49.865	-0.0027	-0.122				0.0500	0	0.00	0.0	0.0	0.0	0.0	-0.022	-0.099	0.000	0.000000
		7.20	16.0	1624.2	100.89	100.88	123.73	-80.4	-80.7	3.88E-07	49.837	-0.081	49.837	-0.0033	-0.093				0.0500	0	0.00	0.0	0.0	0.0	0.0	-0.027	-0.065	0.000	0.000000
		7.26	18.0	1614.7	100.86	100.86	123.71	-71.5	-71.7	3.90E-07	49.846	-0.058	49.846	-0.0031	-0.067				0.0500	1	400.00	7200.0	-26.8	-482.3	129600.0	-0.032	-0.035	-13.825	0.477809
		7.30	20.0	1595.7	100.82	100.82	123.67	-64.3	-64.5	3.95E-07	49.860	-0.028	49.860	-0.0028	-0.032				0.0500	1	400.00	8000.0	-13.0	-259.8	160000.0	-0.037	0.005	2.010	0.010100
		7.34	22.0	1590.2	100.77	100.77	123.62	-58.5	-58.7	3.97E-07	49.880	-0.019	49.880	-0.0024	-0.021				0.0700	1	204.08	4489.8	-4.4	-96.1	98775.5	-0.043	0.021	4.317	0.091321
Calibration freq. >>>		7.38	24.0	1576.0	100.66	100.66	123.50	-53.7	-53.8	4.01E-07	49.926	0.000	49.926	-0.0015	0.000				0.0500	0	0.00	0.0	0.0	0.0	0.0	-0.048	0.047	0.000	0.000000
		7.41	26.0	1562.6	100.49	100.49	123.34	-49.6	-49.6	4.05E-07	49.992	0.015	49.992	-0.0002	0.017				0.0500	0	0.00	0.0	0.0	0.0	0.0	-0.053	0.070	0.000	0.000000
		7.45	28.0	1555.5	100.31	100.31	123.16	-46.2	-46.1	4.08E-07	50.064	0.023	50.064	0.0013	0.026				0.0500	0	0.00	0.0	0.0	0.0	0.0	-0.058	0.084	0.000	0.000000
		7.48	30.0	1546.9	100.08	100.08	122.93	-43.2	-43.0	4.11E-07	50.157	0.032	50.157	0.0031	0.036				0.0500	0	0.00	0.0	0.0	0.0	0.0	-0.063	0.099	0.000	0.000000
phase error vs. f									Sums >>>>		17	6.60E+03	35286.4	-2.5	-645.0	437964.7		-6.59E-02	0.00E+00	2.416419
									Parameter		esd					Chi^2/n-2 >>>>>>>>>>>			0.161095
									D	1.65E+09					esd of an obs of unit weight >>>			0.401
									a	0.013153	0.006545	<<< phase intercept at 0Hz
									b	-0.002532	0.000804	<<< gradient [degrees / MHz]
									tp	0.0070	0.002232	ns <<< time delay

									Phase error, obs-calc

Magnitude error vs f