## Sheet1
FT50-61, 12 turns 0.9mm wire. Lsec = 9.25uH @ 1.5915MHz. 500pF ref. cap. Shield earthed on gen. side. 75 Ohm load.					D. W. Knight. June 2007.				Version	1.00					Ci computed from Rv
		esd					Show columns N-Q to see deatails of fitting procedure.					See sheet 2 for C1'm vs frequency
N =	12.00					Li' = N R0 C2 Rv
R0 =	75.52	0.06	< BNC type, Radiall Model R404012
C2/pF	10.30	0.20
L2 / nH	50.00	20.00	< estimated					from fit	from fit
L2 fit	1	< 1 to include L2 in fit, 0 to exclude it.					Ci =	Li =
1/(N R0 C2)	1.07E+08					-4.831	8.992
			Log (f)	f	x	Rv	esd Rv	Li'	Corr(L2)	y	esd y	fit	weight	w x	w x^2	w y	w x y	y calc	obs-calc	w e	w e^2	Li' calc	Li'					Rv Calc	obs-calc	Computed	Computed
				/MHz	-(2 pi f)^2	Ohms	Ohms	/ uH					w					e			/ uH	obs-calc					Li /uH	Ci /pF
			6.20	1.6	-100.000	959.1	0.77	8.9525	-5.8	111695	89.36	1	1.25E-04	-0.012523019	1.252303	13.987561	-1398.757128588	111692	3.2	0.000399	0.00	8.9528	-0.0003	<< same meas freq as Lsec					959.1	0.0	8.991	-4.920
			6.23	1.7	-114.093	959.4	0.77	8.9553	-6.6	111659	89.33	1	1.25E-04	-0.014296770	1.631156	13.991835	-1596.365203540	111760	-100.6	-0.012610	1.27	8.9472	0.0081					958.5	0.9	9.000	-4.006
			6.26	1.8	-127.910	959.2	0.77	8.9534	-7.4	111682	89.35	1	1.25E-04	-0.016021531	2.049315	13.988819	-1789.310809384	111827	-144.9	-0.018149	2.63	8.9418	0.0116					958.0	1.2	9.003	-3.756
			6.28	1.9	-142.517	958.0	0.77	8.9422	-8.2	111821	89.46	1	1.25E-04	-0.017806513	2.537732	13.971213	-1991.136570562	111897	-76.4	-0.009546	0.73	8.9361	0.0061					957.3	0.7	8.998	-4.352
			6.30	2.0	-157.914	956.2	0.76	8.9254	-9.1	112030	89.63	1	1.24E-04	-0.019656135	3.103972	13.944852	-2202.082698820	111971	58.8	0.007322	0.43	8.9301	-0.0047					956.7	-0.5	8.987	-5.261
			6.34	2.2	-191.076	955.5	0.76	8.9189	-11.0	112110	89.70	1	1.24E-04	-0.023749114	4.537875	13.934405	-2662.523989204	112132	-21.2	-0.002637	0.06	8.9172	0.0017					955.3	0.2	8.993	-4.778
			6.38	2.4	-227.396	953.7	0.76	8.9021	-13.2	112320	89.87	1	1.24E-04	-0.028157022	6.402785	13.907895	-3162.595277961	112307	12.8	0.001588	0.02	8.9031	-0.0010					953.8	-0.1	8.990	-4.945
			6.41	2.6	-266.874	950.9	0.76	8.8760	-15.5	112648	90.13	1	1.23E-04	-0.032851640	8.767252	13.866780	-3700.684528914	112498	150.6	0.018534	2.79	8.8878	-0.0119	< knifing vane					952.2	-1.3	8.979	-5.453
			6.45	2.8	-309.511	950.0	0.76	8.8676	-18.0	112753	90.22	1	1.23E-04	-0.038028040	11.770089	13.853351	-4287.761807297	112704	48.8	0.006000	0.29	8.8714	-0.0038					950.4	-0.4	8.987	-5.047
			6.48	3.0	-355.306	947.7	0.76	8.8461	-20.7	113024	90.44	1	1.22E-04	-0.043443513	15.435730	13.819486	-4910.142874037	112925	98.6	0.012054	1.19	8.8538	-0.0077					948.5	-0.8	8.982	-5.166
			6.54	3.5	-483.611	943.0	0.75	8.8022	-28.3	113579	90.89	1	1.21E-04	-0.058546393	28.313657	13.750041	-6649.665778272	113545	34.6	0.004187	0.14	8.8049	-0.0027					943.3	-0.3	8.987	-4.961
			6.60	4.0	-631.655	937.1	0.75	8.7471	-37.2	114286	91.46	1	1.20E-04	-0.075514878	47.699326	13.662970	-8630.278991287	114260	25.8	0.003084	0.08	8.7491	-0.0020					937.3	-0.2	8.987	-4.931
			6.65	4.5	-799.438	930.3	0.74	8.6837	-47.4	115111	92.13	1	1.18E-04	-0.094191505	75.300264	13.562653	-10842.499800905	115071	40.7	0.004795	0.20	8.6867	-0.0031					930.6	-0.3	8.985	-4.941
			6.70	5.0	-986.960	924.3	0.74	8.6277	-58.9	115847	92.72	1	1.16E-04	-0.114790668	113.293848	13.473879	-13298.185144996	115976	-129.1	-0.015021	1.94	8.6181	0.0096					923.3	1.0	8.998	-4.760
			6.78	6.0	-1421.223	909.5	0.73	8.4895	-86.2	117706	94.23	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	118074	-368.2	0.000000	0.00	8.4631	0.0265					906.7	2.8	9.015	-4.632
			6.85	7.0	-1934.442	893.3	0.71	8.3383	-119.5	119809	95.94	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	120553	-744.5	0.000000	0.00	8.2869	0.0514					887.8	5.5	9.043	-4.508
			6.90	8.0	-2526.619	878.0	0.70	8.1955	-158.8	121860	97.61	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	123414	-1554.6	0.000000	0.00	8.0924	0.1031					867.0	11.0	9.106	-4.278	Li computed from Rv vs Log(f)
			6.95	9.0	-3197.752	859.9	0.69	8.0265	-205.2	124382	99.67	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	126656	-2274.8	0.000000	0.00	7.8826	0.1439					844.5	15.4	9.163	-4.184
			7.00	10.0	-3947.842	842.5	0.67	7.8641	-258.5	126901	101.73	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	130280	-3378.6	0.000000	0.00	7.6606	0.2035					820.7	21.8	9.252	-4.040
			7.08	12.0	-5684.892	803.1	0.64	7.4964	-390.6	133008	106.72	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	138671	-5663.5	0.000000	0.00	7.1911	0.3053					770.4	32.7	9.440	-3.903
			7.15	14.0	-7737.770	763.9	0.61	7.1305	-558.9	139685	112.19	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	148588	-8903.4	0.000000	0.00	6.7048	0.4257					718.3	45.6	9.722	-3.752
			7.20	16.0	-10106.475	725.4	0.58	6.7711	-768.7	146918	118.15	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	160031	-13112.6	0.000000	0.00	6.2189	0.5522					666.2	59.2	10.115	-3.609
			7.26	18.0	-12791.007	686.9	0.55	6.4117	-1027.4	154937	124.77	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	172999	-18062.0	0.000000	0.00	5.7462	0.6655					615.6	71.3	10.619	-3.499
			7.30	20.0	-15791.367	652.4	0.52	6.0897	-1335.5	162877	131.37	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	187493	-24616.5	0.000000	0.00	5.2958	0.7939					567.4	85.0	11.373	-3.356
			7.34	22.0	-19107.554	620.7	0.50	5.7938	-1698.4	170900	138.08	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	203513	-32612.7	0.000000	0.00	4.8730	0.9208					522.1	98.6	12.454	-3.213
			7.38	24.0	-22739.569	592.6	0.47	5.5315	-2117.1	178666	144.63	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	221059	-42392.5	0.000000	0.00	4.4808	1.0507					480.0	112.6	14.098	-3.060
			7.41	26.0	-26687.410	566.5	0.45	5.2879	-2599.2	186513	151.29	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	240130	-53616.6	0.000000	0.00	4.1198	1.1680					441.4	125.1	16.614	-2.919
			7.45	28.0	-30951.079	543.8	0.44	5.0760	-3140.2	193866	157.61	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	260727	-66860.4	0.000000	0.00	3.7898	1.2862					406.0	137.8	21.058	-2.772
			7.48	30.0	-35530.576	527.9	0.42	4.9276	-3713.4	199227	162.35	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	282849	-83622.4	0.000000	0.00	3.4896	1.4379					373.9	154.0	31.949	-2.582
Derived parameters		esd +/-	sums >>>>>>>>>>>>					14	1.72E-03	-0.589576741	322.095305	193.715739	-67121.990603769			-2.82E-14	11.77
D	2.05E-01
a	111208.62	39.23			y obs - y calc
Li / uH	8.992108	0.0032
b = Ci / pF	-4.830782	0.0906
chi^2/(n-2)	0.980624
esd of fit	0.99
Rv0	963.34
				Incremental error analysis
X-over /MHz	Err:502	< complex when Ci is -ve					R0	49.83	49.89	49.95	err contrib
				Li	8.959693	8.992108	8.981270	0.010789
				Ci	6.297673	-4.830782	6.282543	-0.007565

				C2	10.1	10.3	10.5
				Li	8.796297	8.992108	9.144665	0.174184
				Ci	6.414655	-4.830782	6.170287	-0.122184

				Accuracy
				esd Li	0.174547
				esd Ci	0.152275

				Li' vs Log(f)

## Sheet2
Variation of C1'm (nominal capacitance, measured @ 1.5915MHz) with frequency					To see details of fitting calculation show columns N-P
			Log(f)	f	Dial	Ci'm	esd C1'm	y	esd y	fit	weight	x	wx	wy	wxy	wx^2	y calc	e	we	we^2		<Ci'm>	obs-calc
				/MHz		/pF		1/C1'm			w					obs-calc
see sheet 4 for capcal			6.20	1.59	9.33	170.9	0.08	0.005851	0.000003	1	1.33E+11	0	9534947.1	780109694.17	55787.86	6.82E+02	0.005859	-0.000008	-1056209.1	8.367		170.68	0.231
			6.23	1.70	9.33	170.8	0.08	0.005855	0.000003	1	1.33E+11	14.09	1873964830380.6	778542129.04	10971703732.24	2.64E+13	0.005860	-0.000005	-661057.8	3.286		170.65	0.145
			6.26	1.80	9.33	170.8	0.08	0.005855	0.000003	1	1.33E+11	27.91	3711337548690.8	778542129.04	21729167684.7	1.04E+14	0.005861	-0.000006	-788028.8	4.670		170.63	0.173
			6.28	1.90	9.33	170.8	0.08	0.005855	0.000003	1	1.33E+11	42.52	5653702993761.5	778542129.04	33101343863.01	2.40E+14	0.005862	-0.000007	-922255.2	6.396		170.60	0.202
			6.30	2.00	9.32	170.7	0.08	0.005859	0.000003	1	1.33E+11	57.91	7680423036081.7	776976788.75	44997577665.37	4.45E+14	0.005863	-0.000004	-539806.0	2.197		170.57	0.119
			6.34	2.20	9.32	170.7	0.08	0.005859	0.000003	1	1.33E+11	91.08	12078299987780.7	776976788.75	70763581538.75	1.10E+15	0.005865	-0.000006	-843719.7	5.368		170.50	0.185
			6.38	2.40	9.32	170.7	0.08	0.005859	0.000003	1	1.33E+11	127.4	16895022363451.1	776976788.75	98983490542.92	2.15E+15	0.005868	-0.000009	-1176577.4	10.438		170.43	0.258
			6.41	2.60	9.32	170.6	0.08	0.005863	0.000003	1	1.32E+11	166.87	22071247161314.0	775413671.58	129396460903.46	3.68E+15	0.005870	-0.000008	-1013914.3	7.773		170.35	0.223
			6.45	2.80	9.31	170.3	0.08	0.005871	0.000003	1	1.32E+11	209.51	27561954831513.3	772294099.81	161803950066.82	5.77E+15	0.005873	-0.000003	-359022.7	0.980		170.26	0.079
			6.48	3.00	9.30	170.2	0.08	0.005875	0.000003	1	1.31E+11	255.31	33496237582510.0	770737641.78	196773758193.08	8.55E+15	0.005876	-0.000002	-255139.1	0.496		170.17	0.056
			6.54	3.50	9.29	170.0	0.08	0.005882	0.000003	1	1.30E+11	383.61	50059621638902.1	767631372.9	294471543552.96	1.92E+16	0.005885	-0.000003	-378167.3	1.096		169.91	0.084
			6.60	4.00	9.28	169.8	0.08	0.005890	0.000003	1	1.30E+11	531.65	69005759266648.4	764533955.52	406468056746.09	3.67E+16	0.005896	-0.000005	-674306.2	3.503		169.62	0.150
			6.65	4.50	9.26	169.2	0.08	0.005910	0.000003	1	1.28E+11	699.44	89565257227195.2	756829047.64	529354962488.89	6.26E+16	0.005907	0.000003	401022.3	1.256		169.29	-0.090
			6.70	5.00	9.24	168.9	0.08	0.005922	0.000003	1	1.27E+11	886.96	112659301248697.0	752232527.5	667200493651.24	9.99E+16	0.005920	0.000002	277764.4	0.607		168.92	-0.062
			6.78	6.00	9.20	167.9	0.08	0.005955	0.000003	1	1.24E+11	1321.22	164210580216070.0	740071629.43	977799683436.69	2.17E+17	0.005950	0.000004	551788.2	2.450		168.06	-0.125
			6.85	7.00	9.15	166.8	0.08	0.005995	0.000003	1	1.21E+11	1834.44	221854430980886.0	725066793.5	1330093314234.23	4.07E+17	0.005986	0.000010	1181443.1	11.541		167.07	-0.272
			6.90	8.00	9.10	165.7	0.08	0.006037	0.000003	1	1.18E+11	2426.62	285517715605494.0	710278457.82	1723575006892.53	6.93E+17	0.006026	0.000010	1196610.2	12.170		165.93	-0.280
			6.95	9.00	9.04	164.3	0.08	0.006087	0.000003	1	1.14E+11	3097.75	352584876631061.0	692815827.17	2146171493662.11	1.09E+18	0.006073	0.000014	1603718.6	22.596		164.67	-0.381
			7.00	10.00	8.99	163.1	0.08	0.006129	0.000003	1	1.11E+11	3847.84	425933464019376.0	678497771.34	2610752058928.72	1.64E+18	0.006125	0.000005	528510.1	2.523		163.27	-0.127
			7.08	12.00	8.85	160.0	0.08	0.006252	0.000003	1	1.02E+11	5584.89	571326390237989.0	639524413.45	3571674866830.93	3.19E+18	0.006245	0.000007	695823.2	4.733		160.13	-0.174
			7.15	14.00	8.69	156.3	0.08	0.006397	0.000003	1	9.33E+10	7637.77	712778513593204.0	596961098	4559451476213.48	5.44E+18	0.006387	0.000010	944364.5	9.556		156.58	-0.248
			7.20	16.00	8.50	152.0	0.08	0.006577	0.000003	1	8.35E+10	10006.47	835330058231609.0	549082693.5	5494382194221.76	8.36E+18	0.006550	0.000027	2270598.9	61.760		152.66	-0.631
			7.26	18.00	8.34	148.3	0.08	0.006742	0.000004	1	7.56E+10	12691.01	959555321627486.0	509783118.55	6469661280895.76	1.22E+19	0.006736	0.000007	494571.9	3.235		148.46	-0.144
			7.30	20.00	8.14	143.9	0.08	0.006947	0.000004	1	6.71E+10	15691.37	1052476160993050.0	465981442.37	7311885846800.62	1.65E+19	0.006943	0.000004	279437.9	1.164		144.03	-0.086
			7.34	22.00	7.94	139.5	0.08	0.007170	0.000004	1	5.91E+10	19007.55	1123860376276350.0	423930437.63	8057880736565.21	2.14E+19	0.007172	-0.000002	-147306.1	0.367		139.42	0.048
			7.38	24.00	7.73	134.7	0.08	0.007424	0.000004	1	5.14E+10	22639.57	1164433442017410.0	381847454.08	8644861608493.71	2.64E+19	0.007423	0.000001	40555.5	0.032		134.71	-0.014
			7.41	26.00	7.51	129.9	0.08	0.007695	0.000005	1	4.46E+10	26587.41	1184595210122640.0	342866818.36	9115940778210.74	3.15E+19	0.007696	-0.000001	-31740.9	0.023		129.94	0.012
			7.45	28.00	7.31	125.4	0.08	0.007971	0.000005	1	3.87E+10	30851.08	1193816154664930.0	308464594.11	9516465685655.72	3.68E+19	0.007991	-0.000019	-746475.6	14.400		125.14	0.303
			7.48	30.00	7.09	120.8	0.08	0.008281	0.000005	1	3.32E+10	35430.58	1177266213050670.0	275154967.97	9748898961459.84	4.17E+19	0.008307	-0.000026	-872482.5	22.910		120.38	0.382
			sums >>>>>>>					29	3.08E+12		11877851047520100.0	18846666281.52	83945511138919.4	2.08E+20			-4.13E-06	225.894
Derived parameters		esd					Chi^2(/n-2) >>>>>		8.366
Denom	4.98E+32					esd of fit >>>>>>		2.892
a	0.005859	0.000002
C1' /pF	170.683	0.054					Ci'measured vs Log(f) (with expectation value)
b	6.91E-08	2.27E-10
L1 /nH	69.105	0.227

KT	18.07	< Xformer const. = ( N R0 / k Ri) + (1/N) -1
L2 /nH	50	20
L1(true)	71.871	1.130	nH	< corrected for L2

				C1'm obs-calc vs Log(f)

## Sheet3
Stray capacitance of voltage sampling network				Predicted C1 = C1' + C1s = C2 [( N R0 / k Ri) + (1/N) -1]
Parameter		esd +/-	Deriv dC1/dpar	(der*esd)^2
N =	12
R0 =	75.52	0.06	2.59	0.02
C2 /pF =	10.3	0.2	18.07	13.07
Ri =	49.76	0.06	-3.93	0.06
k =	0.959	0.0089	-203.97	3.3
				16.44	<<< sum, used to calculate esd of predicted C1.
KT	18.074	< transformer const = ( N R0 / k Ri) + (1/N) -1

C1' =	170.683	1.708
C1 pred =	186.164	4.055
C1s =	15.481	4.400

C1 Prediction for bridge evaluation (see sheet 5)
R0 =	75.5
C1 pred =	186.11

Self capacitance of transformer secondary				Ci' = Ci apparent + ( Csh * Ro / Ri ) + [ L2 / ( N Ro Rv ) ] + ( LH / Ri^2 ) - Cis - [ L0 / ( R0 Ri ) ] + C0 Ro / Ri
parameter		esd +/-
Csh / pF	0.00	0.00	< shield protrusion capacitance. Set to 0 for gen. side earth.
L2 / nH	50.00	20.00	< inductance of upper voltage sampling arm
L2 fit	1		< fitting flag =1 when L2 is included in fit on sheet 1, 0 otherwise.
RV0	963.34		< optimal Rv
LH / nH	15.00	2.00	< inductance of sec. load resistor
Cis / pF	0.40	0.10	< strays accross sec winding
L0 / nH	14.00	3.00	< inductance of pri. transmission line
C0 / pF	5.00	0.50	< capacitance of pri. transmission line
Rline	52.92

Series for Ci'		esd +/-	Variance
Ci apparent	-4.83	0.09	0.008	< from fit on sheet 1
Corr for Csh	0.00	0.00	0.000
Corr for L2	0.00	0.02	0.001	< inclusion depends on whether L2 is included in fit on sheet 1.
Corr for LH	6.06	0.81	0.652
Corr for Cis	-0.40	0.10	0.010
Corr for L0	-3.73	0.80	0.637
Corr for C0	7.59	0.76	0.576
Ci' / pF	4.69	1.37	1.88	Only valid if parameters have been filled in correctly!

lw / mm	228.00	< length of secondary winding wire
Ci' pred / pF	7.64	< predicted Ci' if phase velocity = c

## Sheet4
Data reduction for 500pF ref. cap.					C =	69.83	*exp( -	0.32	* x) -	56.91325	+	24.040788	* x
Non linearity at low settings is corrected by subtracting an exponential function f(x) =( C0-a)Exp(-px)					where x is the dial reading
seed for C0	12.92	<<< seed values are required because the fitting process is iterative
seed for -a	56.91	<<<
C0 - a	69.83					Show columns J-M for details of fitting calculation.
p	0.32
		Dial	f(x)	C meas	esd C	y	fit	w	wy	wx	wx^2	wxy	y calc	obs-calc	we	we^2		C calc	obs-calc	w ec	w (ec)^2
		x		/ pF		C - f(x)					e					ec
0 - 25pF range		0	69.83	15.05	1.06	-54.8	1	0.890	-48.74	0.00	0.00	0.00	-56.91	2.13	1.90	4.049		12.92	2.13	1.898	4.049
		1	50.71	17.40	0.87	-33.3	1	1.322	-44.04	1.32	1.32	-44.04	-32.87	-0.43	-0.57	0.250		17.83	-0.43	-0.575	0.250
0 - 250pF range		2	36.82	28.50	0.91	-8.3	1	1.197	-9.96	2.39	4.79	-19.92	-8.83	0.51	0.61	0.312		27.99	0.51	0.611	0.312
		3	26.74	41.50	0.97	14.8	1	1.072	15.82	3.21	9.64	47.46	15.21	-0.45	-0.48	0.214		41.95	-0.45	-0.479	0.214
		4	19.42	58.00	1.03	38.6	1	0.939	36.23	3.76	15.02	144.92	39.25	-0.67	-0.62	0.416		58.67	-0.67	-0.625	0.416
		5	14.10	76.00	1.10	61.9	1	0.820	50.79	4.10	20.51	253.94	63.29	-1.39	-1.14	1.583		77.39	-1.39	-1.140	1.583
		6	10.24	96.00	1.18	85.8	1	0.713	61.18	4.28	25.68	367.07	87.33	-1.57	-1.12	1.756		97.57	-1.57	-1.119	1.756
		7	7.43	118.00	1.27	110.6	1	0.618	68.34	4.33	30.28	478.35	111.37	-0.81	-0.50	0.402		118.81	-0.81	-0.498	0.402
		8	5.40	141.00	1.36	135.6	1	0.537	72.88	4.30	34.40	583.08	135.41	0.19	0.10	0.019		140.81	0.19	0.101	0.019
		9	3.92	163.00	1.45	159.1	1	0.474	75.45	4.27	38.42	679.09	159.45	-0.37	-0.18	0.066		163.37	-0.37	-0.177	0.066
		10	2.85	187.00	1.55	184.2	1	0.417	76.85	4.17	41.73	768.49	183.49	0.66	0.27	0.181		186.34	0.66	0.275	0.181
		11	2.07	212.00	1.65	209.9	1	0.368	77.30	4.05	44.55	850.27	207.54	2.40	0.88	2.117		209.60	2.40	0.883	2.117
		12	1.50	236.50	1.75	235.0	1	0.328	77.09	3.94	47.24	925.04	231.58	3.42	1.12	3.843		233.08	3.42	1.123	3.843
0 - 2500pF range		13	1.09	260.00	3.54	258.9	1	0.080	20.66	1.04	13.49	268.59	255.62	3.29	0.26	0.865		256.71	3.29	0.263	0.865
		14	0.79	280.00	3.62	279.2	1	0.076	21.31	1.07	14.96	298.29	279.66	-0.45	-0.03	0.015		280.45	-0.45	-0.034	0.015
		15	0.57	302.00	3.71	301.4	1	0.073	21.92	1.09	16.36	328.85	303.70	-2.27	-0.17	0.376		304.27	-2.27	-0.165	0.376
		16	0.42	327.00	3.81	326.6	1	0.069	22.52	1.10	17.65	360.35	327.74	-1.16	-0.08	0.092		328.16	-1.16	-0.080	0.092
		17	0.30	350.00	3.90	349.7	1	0.066	22.99	1.12	19.00	390.85	351.78	-2.08	-0.14	0.285		352.08	-2.08	-0.137	0.285
		18	0.22	375.00	4.00	374.8	1	0.063	23.42	1.13	20.25	421.63	375.82	-1.04	-0.07	0.068		376.04	-1.04	-0.065	0.068
		19	0.16	400.00	4.10	399.8	1	0.059	23.79	1.13	21.48	451.93	399.86	-0.02	0.00	0.000		400.02	-0.02	-0.001	0.000
		20	0.12	425.00	4.20	424.9	1	0.057	24.09	1.13	22.68	481.73	423.90	0.98	0.06	0.055		424.02	0.98	0.056	0.055
		21	0.08	448.00	4.29	447.9	1	0.054	24.32	1.14	23.94	510.62	447.94	-0.03	0.00	0.000		448.03	-0.03	-0.001	0.000
		22	0.06	470.00	4.38	469.9	1	0.052	24.50	1.15	25.23	538.91	471.98	-2.05	-0.11	0.218		472.05	-2.05	-0.107	0.218
		23	0.04	498.00	4.49	498.0	1	0.050	24.68	1.14	26.22	567.60	496.02	1.93	0.10	0.185		496.07	1.93	0.096	0.185
		24	0.03	520.00	4.58	520.0	1	0.048	24.79	1.14	27.46	594.92	520.07	-0.10	0.00	0.000		520.10	-0.10	-0.005	0.000
		25	0.02	542.00	4.67	542.0	1	0.046	24.87	1.15	28.68	621.81	544.11	-2.13	-0.10	0.208		544.13	-2.13	-0.098	0.208
		sums >>>>>>>>>					26	10.489	813.03	58.65	590.98	10869.79			0.00	17.576				0.000	17.576
D	2758.883
a	-56.913247
esd a	0.156871
b	24.040788
esd b	0.002784
Chi^2/n-2)	0.732319
esd fit	0.855756

## Sheet5
Bridge performance evaluation - change Rvset and Cvset to see the effect of different calibration choices.
N	12
R0	75.5
C2 /pF	10.3
Li /uH	8.992
Rv0	963.6	<<< Guide value (estimated value for zero phase error at 0Hz).
Rvset	960.00	<<< selected value (usually set to balance bridge at some chosen 'calibration' frequency, but can be arbitrary)
C1s /pF	15.481
L1 /nH	69.105
C1 pred	186.112					Show columns W-Z for details of fit.
Cv set 0	196.412	<<< Guide value (estimated value for zero magnitude error to zero at 0Hz)
C1 set	189.200
Cv set	199.500	<<< selected value (see note for Rvset)
			x					y				Determining the gradient of Phi vs. f.
			f	Rv	C1'm	C1'0	Cv	Xcv	Xcvset	U	Rbal	Xbal	|Zbal|	(|Zbal|/R0)-1	Phi				esd Phi	fit	weight					Phi calc	obs-calc
		Log(f)	/MHz	Ohms	/pF	/pF	/pF					Magn. err.	deg.				esd y		w	wx	wy	wxy	wx^2	y calc	e	we	we^2
		6.20	1.59	959.1	170.91	170.71	196.49	-508.92	-501.25	8.60E-07	76.39	0.51	76.39	0.012	0.381				0.0500	1	400	637	152	242	1013	0.359	0.022	8.8	0.193
		6.23	1.70	959.4	170.80	170.60	196.38	-476.73	-469.28	8.85E-07	76.45	0.50	76.45	0.013	0.372				0.0200	1	2500	4250	931	1583	7225	0.367	0.006	13.9	0.078
		6.26	1.80	959.2	170.80	170.60	196.38	-450.25	-443.21	9.04E-07	76.47	0.49	76.47	0.013	0.364				0.0075	1	17778	32000	6470	11645	57600	0.374	-0.010	-184.2	1.909
		6.28	1.90	958.0	170.80	170.60	196.38	-426.55	-419.88	9.22E-07	76.48	0.50	76.48	0.013	0.378				0.0075	1	17778	33778	6717	12762	64178	0.382	-0.004	-69.1	0.269
		6.30	2.00	956.2	170.69	170.48	196.26	-405.46	-398.88	9.39E-07	76.51	0.55	76.51	0.013	0.415				0.0075	1	17778	35556	7385	14769	71111	0.389	0.026	467.0	12.268
		6.34	2.20	955.5	170.69	170.48	196.26	-368.60	-362.62	9.65E-07	76.54	0.54	76.54	0.014	0.402				0.0075	1	17778	39111	7143	15714	86044	0.404	-0.002	-38.6	0.084
		6.38	2.40	953.7	170.69	170.48	196.26	-337.88	-332.40	9.86E-07	76.55	0.55	76.55	0.014	0.410				0.0075	1	17778	42667	7288	17490	102400	0.419	-0.009	-157.8	1.401
		6.41	2.60	950.9	170.57	170.37	196.15	-312.07	-306.83	1.01E-06	76.59	0.59	76.60	0.015	0.444				0.0075	1	17778	46222	7899	20538	120178	0.434	0.011	189.8	2.026
		6.45	2.80	950.0	170.34	170.14	195.92	-290.12	-284.92	1.02E-06	76.69	0.60	76.70	0.016	0.452				0.0075	1	17778	49778	8033	22491	139378	0.448	0.003	59.5	0.199
		6.48	3.00	947.7	170.23	170.03	195.81	-270.94	-265.92	1.04E-06	76.74	0.63	76.75	0.016	0.472				0.0075	1	17778	53333	8387	25161	160000	0.463	0.008	149.8	1.262
		6.54	3.50	943.0	170.00	169.80	195.58	-232.50	-227.93	1.06E-06	76.85	0.66	76.85	0.018	0.494				0.0075	1	17778	62222	8781	30734	217778	0.500	-0.007	-115.8	0.755
		6.60	4.00	937.1	169.77	169.57	195.35	-203.68	-199.44	1.08E-06	76.95	0.71	76.95	0.019	0.526				0.0075	1	17778	71111	9353	37412	284444	0.538	-0.011	-203.7	2.334
		6.65	4.50	930.3	169.20	169.00	194.78	-181.58	-177.28	1.10E-06	77.18	0.78	77.18	0.022	0.581				0.0075	1	17778	80000	10336	46511	360000	0.575	0.007	119.5	0.804
		6.70	5.00	924.3	168.85	168.66	194.44	-163.71	-159.55	1.12E-06	77.32	0.82	77.32	0.024	0.607				0.0075	1	17778	88889	10798	53991	444444	0.612	-0.004	-77.8	0.340
		6.78	6.00	909.5	167.94	167.74	193.52	-137.07	-132.96	1.15E-06	77.69	0.93	77.70	0.029	0.685				0.0075	1	17778	106667	12176	73057	640000	0.686	-0.001	-19.4	0.021
		6.85	7.00	893.3	166.80	166.60	192.38	-118.18	-113.97	1.19E-06	78.15	1.05	78.16	0.035	0.766				0.0075	1	17778	124444	13620	95342	871111	0.760	0.006	105.3	0.624
		6.90	8.00	878.0	165.65	165.46	191.25	-104.03	-99.72	1.22E-06	78.62	1.13	78.63	0.041	0.826				0.0075	1	17778	142222	14693	117545	1137778	0.834	-0.008	-141.4	1.124
		6.95	9.00	859.9	164.29	164.10	189.88	-93.13	-88.64	1.26E-06	79.19	1.25	79.20	0.049	0.906				0.0075	1	17778	160000	16106	144951	1440000	0.909	-0.003	-48.4	0.132
		7.00	10.00	842.5	163.15	162.96	188.74	-84.32	-79.78	1.29E-06	79.67	1.34	79.68	0.055	0.965				0.0075	1	17778	177778	17158	171579	1777778	0.983	-0.018	-315.5	5.600
		7.08	12.00	803.1	159.96	159.78	185.56	-71.47	-66.48	1.38E-06	81.03	1.59	81.04	0.073	1.124				0.0075	1	17778	213333	19987	239848	2560000	1.131	-0.007	-125.1	0.880
		7.15	14.00	763.9	156.33	156.16	181.94	-62.48	-56.98	1.49E-06	82.64	1.85	82.66	0.095	1.279				0.0075	1	17778	248889	22739	318351	3484444	1.280	-0.001	-12.0	0.008
		7.20	16.00	725.4	152.03	151.87	177.65	-55.99	-49.86	1.60E-06	84.62	2.13	84.65	0.121	1.441				0.0075	1	17778	284444	25611	409769	4551111	1.428	0.012	220.2	2.728
		7.26	18.00	686.9	148.32	148.16	173.95	-50.83	-44.32	1.73E-06	86.41	2.40	86.45	0.145	1.589				0.0075	1	17778	320000	28248	508469	5760000	1.577	0.012	219.0	2.697
		7.30	20.00	652.4	143.94	143.80	169.58	-46.93	-39.89	1.87E-06	88.63	2.68	88.67	0.174	1.735				0.0075	1	17778	355556	30842	616841	7111111	1.725	0.010	173.8	1.699
		7.34	22.00	620.7	139.47	139.34	165.12	-43.81	-36.26	2.02E-06	91.01	2.98	91.06	0.206	1.874				0.0075	1	17778	391111	33322	733078	8604444	1.874	0.001	14.5	0.012
		7.38	24.00	592.6	134.70	134.57	160.35	-41.36	-33.24	2.18E-06	93.70	3.29	93.76	0.242	2.009				0.0075	1	17778	426667	35714	857137	10240000	2.022	-0.013	-232.2	3.032
		7.41	26.00	566.5	129.95	129.83	155.61	-39.34	-30.68	2.35E-06	96.54	3.61	96.61	0.280	2.142				0.0075	0	0	0	0	0	0	2.170	-0.029	0.0	0.000
		7.45	28.00	543.8	125.45	125.34	151.12	-37.61	-28.49	2.52E-06	99.40	3.92	99.48	0.318	2.257				0.0075	0	0	0	0	0	0	2.319	-0.062	0.0	0.000
		7.48	30.00	527.9	120.76	120.66	146.44	-36.23	-26.59	2.68E-06	102.57	4.19	102.65	0.360	2.339				0.0075	0	0	0	0	0	0	2.467	-0.128	0.0	0.000
					Sums >>>>		26	429567	3590664.4	369887.76	4597010.39	50293571.55			0	42.48
	Phase error vs f					Parameter		+/- esd					Chi^2/(n-2) >>>>		1.769832
					D	8.71E+12					esd of an obs of unit w >>>>			1.330
					a	0.240675	0.003196	phase intercept at 0Hz
					b	0.074221	0.000295	degrees / MHz
					tp	-0.206169	0.000821	ns (propagation delay)

					Phase err. Deviation from linearity vs. Log(f)

		Magnitude error vs f