## Sheet1
FT50-61, 12 turns 0.9mm wire + 2 turn quadrature winding with 250 Ohm pot. Lsec = 9.25uH @ 1.5915MHz. 40pF variable ref cap with padding. No shield. L-bal coil.					D. W. Knight. June 2007.			Version	1.01
		esd +/-					Show columns N-Q to see deatails of fitting procedure.					See sheet 2 for C1'm vs frequency
N =	12.00					Li' = N R0 C2 Rv
R0 =	49.89	0.06					from fit	from fit
C2+Cx/pF	10.90	0.20					Ci =	Li =
L2 / nH	900.00	20.00					Li computed from Rv
L2 fit	1	< 1 to include L2 in fit, 0 to exclude it.
1/(N R0 C2)	1.53E+08					0.469	9.993
Rq	40.17	< DC resistance of the quadrature network.
			Log (f)	f	x	Rv + Rq	Rv	esd Rv	Li'	Corr(L2)	y	esd y	fit	weight	w x	w x^2	w y	w x y	y calc	obs-calc	w e	w e^2	Li' calc	Li'					Computed	Computed
				/MHz	-(2 pi f)^2		Ohms	Ohms	/ uH					w					e			/ uH	obs-calc					Li /uH	Ci /pF
			6.20	1.59	-100.000	1570.9	1530.7	15.31	9.9890	-98.2	100012.41	1001.1	1	9.98E-07	-0.000099779	0.009978	0.099792	-9.979158026	100018.52	-6.11	-0.000006	0.00	9.9883	0.0006					9.984	-0.452
			6.23	1.70	-114.093	1570.2	1530.0	7.65	9.9844	-112.1	100044.32	500.8	1	3.99E-06	-0.000454946	0.051906	0.398928	-45.514775957	100011.90	32.42	0.000129	0.00	9.9876	-0.0032					9.979	-0.797
			6.26	1.80	-127.910	1569.9	1529.7	0.92	9.9824	-125.7	100050.36	60.1	1	2.77E-04	-0.035405792	4.528757	27.694163	-3542.362355340	100005.42	44.95	0.012441	0.56	9.9869	-0.0045					9.976	-0.865
			6.28	1.90	-142.517	1569.9	1529.7	0.92	9.9824	-140.1	100036.01	60.1	1	2.77E-04	-0.039449046	5.622163	27.690189	-3946.325110118	99998.56	37.45	0.010366	0.39	9.9862	-0.0037					9.976	-0.776
Calibration freq. >>>			6.30	2.00	-157.914	1569.9	1529.7	0.92	9.9824	-155.2	100020.88	60.1	1	2.77E-04	-0.043710854	6.902541	27.686001	-4371.998028566	99991.33	29.55	0.008179	0.24	9.9854	-0.0029					9.975	-0.700
			6.34	2.20	-191.076	1568.5	1528.3	0.92	9.9733	-187.9	100079.88	60.2	1	2.76E-04	-0.052793368	10.087521	27.651650	-5283.554066775	99975.76	104.12	0.028768	3.00	9.9837	-0.0104					9.964	-1.059
			6.38	2.40	-227.396	1568.2	1528.0	0.92	9.9713	-223.7	100063.80	60.2	1	2.76E-04	-0.062803808	14.281315	27.636354	-6284.387597445	99958.71	105.09	0.029024	3.05	9.9818	-0.0105					9.961	-0.976
			6.41	2.60	-266.874	1572.3	1532.1	0.92	9.9981	-261.9	99757.29	60.0	1	2.78E-04	-0.074103319	19.776257	27.699752	-7392.346361131	99940.17	-182.88	-0.050781	9.29	9.9798	0.0182					9.986	0.174
			6.45	2.80	-309.511	1571.1	1530.9	0.92	9.9903	-303.9	99793.62	60.1	1	2.77E-04	-0.085807740	26.558422	27.666449	-8563.064699512	99920.16	-126.54	-0.035081	4.44	9.9776	0.0126					9.976	-0.104
			6.48	3.00	-355.306	1570.0	1529.8	0.92	9.9831	-349.1	99820.37	60.1	1	2.77E-04	-0.098362280	34.948685	27.634112	-9818.559297637	99898.66	-78.28	-0.021672	1.70	9.9753	0.0078					9.966	-0.293
			6.54	3.50	-483.611	1567.8	1527.6	0.92	9.9687	-475.9	99837.86	60.2	1	2.76E-04	-0.133497206	64.560666	27.559519	-13328.075969832	99838.42	-0.55	-0.000152	0.00	9.9687	0.0001					9.946	-0.513
			6.60	4.00	-631.655	1566.5	1526.3	0.92	9.9602	-622.1	99777.09	60.2	1	2.76E-04	-0.174067060	109.950274	27.495885	-17367.904188887	99768.91	8.18	0.002253	0.02	9.9610	-0.0008					9.931	-0.528
			6.65	4.50	-799.438	1564.7	1524.5	0.91	9.9485	-788.3	99729.45	60.3	1	2.75E-04	-0.219784322	175.703929	27.417973	-21918.968473609	99690.14	39.31	0.010807	0.42	9.9524	-0.0039					9.911	-0.566
			6.70	5.00	-986.960	1563.5	1523.3	0.91	9.9407	-974.0	99622.95	60.4	1	2.74E-04	-0.270911681	267.379112	27.345595	-26989.020581556	99602.09	20.86	0.005725	0.12	9.9427	-0.0021					9.895	-0.538
			6.78	6.00	-1421.223	1561.6	1521.4	0.91	9.9283	-1404.3	99318.27	60.4	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	99398.21	-79.94	0.000000	0.00	9.9204	0.0079					9.863	-0.462
			6.85	7.00	-1934.442	1566.3	1526.1	0.92	9.9589	-1905.5	98506.86	60.2	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	99157.25	-650.39	0.000000	0.00	9.8948	0.0641					9.870	-0.179
			6.90	8.00	-2526.619	1573.4	1533.2	0.92	10.0053	-2477.3	97470.08	60.0	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	98879.23	-1409.15	0.000000	0.00	9.8662	0.1391					9.888	0.047
			6.95	9.00	-3197.752	1576.2	1536.0	0.92	10.0235	-3129.6	96635.57	59.9	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	98564.13	-1928.56	0.000000	0.00	9.8334	0.1901					9.875	0.094	Ci computed from Rv
			7.00	10.00	-3947.842	1580.9	1540.7	0.92	10.0542	-3852.0	95608.91	59.7	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	98211.96	-2603.05	0.000000	0.00	9.7978	0.2564					9.870	0.153
			7.08	12.00	-5684.892	1586.3	1546.1	0.93	10.0894	-5527.4	93586.05	59.5	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	97396.42	-3810.37	0.000000	0.00	9.7159	0.3735					9.825	0.167
			7.15	14.00	-7737.770	1594.8	1554.6	0.93	10.1449	-7482.3	91089.25	59.1	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	96432.59	-5343.34	0.000000	0.00	9.6233	0.5217					9.784	0.193
			7.20	16.00	-10106.475	1600.9	1560.7	0.94	10.1847	-9734.6	88451.68	58.9	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	95320.49	-6868.80	0.000000	0.00	9.5188	0.6659					9.715	0.186
			7.26	18.00	-12791.007	1586.3	1546.1	0.93	10.0894	-12436.7	86676.75	59.5	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	94060.10	-7383.35	0.000000	0.00	9.3900	0.6995					9.513	0.074
			7.30	20.00	-15791.367	1588.8	1548.6	0.93	10.1058	-15329.2	83624.27	59.4	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	92651.43	-9027.16	0.000000	0.00	9.2609	0.8448					9.401	0.070
			7.34	22.00	-19107.554	1587.1	1546.9	0.93	10.0947	-18568.7	80493.50	59.4	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	91094.48	-10600.99	0.000000	0.00	9.1188	0.9758					9.256	0.053
			7.38	24.00	-22739.569	1565.8	1525.6	0.92	9.9557	-22406.8	78038.43	60.3	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	89389.26	-11350.82	0.000000	0.00	8.9449	1.0108					8.999	-0.017
			7.41	26.00	-26687.410	1555.1	1514.9	0.91	9.8858	-26482.7	74672.07	60.7	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	87535.75	-12863.68	0.000000	0.00	8.7705	1.1153					8.796	-0.041
			7.45	28.00	-30951.079	1540.2	1500.0	0.90	9.7886	-31018.7	71140.81	61.3	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	85533.96	-14393.15	0.000000	0.00	8.5798	1.2088					8.570	-0.068
			7.48	30.00	-35530.576	1515.1	1474.9	0.88	9.6248	-36214.2	67683.86	62.3	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	83383.88	-15700.02	0.000000	0.00	8.3613	1.2635					8.293	-0.108
Derived parameters		esd +/-	sums >>>>>>>>>>>>					14	3.32E-03	-1.291251201	740.361526	331.676362	-128862.060664392			-8.52E-14	23.22
D	7.91E-01
a	100065.47	42.56
Li / uH	9.993457	0.0043		y obs - y calc
b = Ci / pF	0.469499	0.0901
chi^2/(n-2)	1.935
esd of fit	1.39
Rv0	1531.42
				Incremental error analysis
X-over /MHz	73.48					R0	49.83	49.89	49.95	err contrib.
				Li	8.994662	9.993457	9.016323	0.010831
				Ci	6.173432	0.469499	6.158601	-0.007416

				C2	10.1	10.3	10.5
				Li	8.828914	9.993457	9.182070	0.176578
				Ci	6.289328	0.469499	6.047431	-0.120949

				Accuracy
				esd Li	0.176961
				esd Ci	0.151019

				Li' vs Log(f)

## Sheet2
Variation of C1'm (nominal capacitance, measured @ 1.5915MHz) with frequency					To see details of fitting calculation show columns N-P
			Log(f)	f	Dial	Ci'm	esd C1'm	y	esd y	fit	weight	x	wx	wy	wxy	wx^2	y calc	e	we	we^2		<Ci'm>	obs-calc
				/MHz		/pF		1/C1'm			w					obs-calc
			6.20	1.59	5.490	100.87	0.63	0.009913	0.000061	1	2.65E+08	0.0	18955	2627623	187.91	1.36E+00	0.009929	-0.000016	-4243.7	0.068		100.71	0.163
			6.23	1.70	5.490	100.87	0.13	0.009913	0.000012	1	6.63E+09	14.1	93383338594	65690577	925752789	1.32E+12	0.009929	-0.000016	-106028.3	1.697		100.71	0.163
			6.26	1.80	5.490	100.87	0.10	0.009913	0.000010	1	1.04E+10	27.9	288973782123	102641526	2864732498	8.07E+12	0.009929	-0.000016	-165571.7	2.648		100.71	0.162
			6.28	1.90	5.490	100.87	0.10	0.009913	0.000010	1	1.04E+10	42.5	440211087155	102641526	4364018764	1.87E+13	0.009929	-0.000016	-165468.5	2.644		100.71	0.162
			6.30	2.00	5.490	100.87	0.08	0.009913	0.000008	1	1.62E+10	57.9	936911533826	160377385	9288043016	5.43E+13	0.009929	-0.000016	-258374.6	4.127		100.71	0.162
			6.34	2.20	5.480	100.85	0.08	0.009916	0.000008	1	1.62E+10	91.1	1471910788567	160256186	14595418865	1.34E+14	0.009929	-0.000013	-217369.3	2.924		100.71	0.137
			6.38	2.40	5.490	100.87	0.08	0.009913	0.000008	1	1.62E+10	127.4	2060972584739	160377385	20431386881	2.63E+14	0.009929	-0.000016	-257607.6	4.102		100.71	0.162
			6.41	2.60	5.490	100.87	0.08	0.009913	0.000008	1	1.62E+10	166.9	2699643636395	160377385	26762832260	4.51E+14	0.009929	-0.000016	-257171.9	4.088		100.71	0.161
			6.45	2.80	5.485	100.86	0.08	0.009915	0.000008	1	1.62E+10	209.5	3387700654609	160316778	33588095415	7.10E+14	0.009929	-0.000015	-236374.6	3.455		100.71	0.148
			6.48	3.00	5.485	100.86	0.08	0.009915	0.000008	1	1.62E+10	255.3	4128185800847	160316778	40929796551	1.05E+15	0.009929	-0.000015	-235869.4	3.441		100.71	0.148
			6.54	3.50	5.475	100.83	0.08	0.009917	0.000008	1	1.62E+10	383.6	6196570851370	160195609	61452736316	2.38E+15	0.009929	-0.000012	-193848.4	2.326		100.71	0.122
			6.60	4.00	5.460	100.80	0.08	0.009921	0.000008	1	1.61E+10	531.7	8574987682545	160013971	85072176852	4.56E+15	0.009929	-0.000008	-131425.9	1.071		100.71	0.083
			6.65	4.50	5.450	100.77	0.08	0.009923	0.000008	1	1.61E+10	699.4	11269767602489	159892955	111835201852	7.88E+15	0.009929	-0.000006	-89130.6	0.493		100.72	0.056
			6.70	5.00	5.435	100.73	0.08	0.009927	0.000008	1	1.61E+10	887.0	14269628247862	159711546	141657823004	1.27E+16	0.009929	-0.000002	-26513.3	0.044		100.72	0.017
			6.78	6.00	5.400	100.64	0.08	0.009936	0.000008	1	1.60E+10	1321.2	21181160415679	159288791	210456019794	2.80E+16	0.009929	0.000007	118999.8	0.883		100.72	-0.075
			6.85	7.00	5.365	100.56	0.08	0.009945	0.000008	1	1.60E+10	1834.4	29304988433636	158866783	291431972597	5.38E+16	0.009928	0.000017	264594.3	4.383		100.72	-0.168
			6.90	8.00	5.390	100.62	0.08	0.009939	0.000008	1	1.60E+10	2426.6	38863006515671	159168141	386240392064	9.43E+16	0.009928	0.000011	171154.9	1.829		100.73	-0.108
			6.95	9.00	5.380	100.59	0.08	0.009941	0.000008	1	1.60E+10	3097.8	49561289682886	159047552	492689845262	1.54E+17	0.009927	0.000014	218484.3	2.984		100.73	-0.138
			7.00	10.00	5.380	100.59	0.08	0.009941	0.000008	1	1.60E+10	3847.8	61562065283973	159047552	611989813276	2.37E+17	0.009927	0.000014	226672.5	3.211		100.74	-0.144
			7.08	12.00	5.370	100.57	0.08	0.009944	0.000008	1	1.60E+10	5584.9	89263062536271	158927024	887590287283	4.99E+17	0.009926	0.000018	285542.1	5.101		100.75	-0.181
			7.15	14.00	5.375	100.58	0.08	0.009942	0.000008	1	1.60E+10	7637.8	122135824021591	158987281	1214308258141	9.33E+17	0.009924	0.000018	287994.3	5.187		100.76	-0.183
			7.20	16.00	5.375	100.58	0.08	0.009942	0.000008	1	1.60E+10	10006.5	160013862974729	158987281	1590902233508	1.60E+18	0.009923	0.000020	313838.7	6.159		100.78	-0.199
			7.26	18.00	5.365	100.56	0.08	0.009945	0.000008	1	1.60E+10	12691.0	202737251142530	158866783	2016179503132	2.57E+18	0.009921	0.000024	382927.8	9.179		100.80	-0.243
			7.30	20.00	5.370	100.57	0.08	0.009944	0.000008	1	1.60E+10	15691.4	250794365166337	158927024	2493782268970	3.94E+18	0.009919	0.000025	395755.5	9.799		100.82	-0.251
			7.34	22.00	5.395	100.63	0.08	0.009937	0.000008	1	1.60E+10	19007.6	304565342044549	159228459	3026543542633	5.79E+18	0.009917	0.000021	332406.8	6.896		100.84	-0.211
			7.38	24.00	5.420	100.69	0.08	0.009931	0.000008	1	1.61E+10	22639.6	363679608785095	159530274	3611696567149	8.23E+18	0.009914	0.000017	272323.5	4.617		100.87	-0.172
			7.41	26.00	5.472	100.83	0.08	0.009918	0.000008	1	1.61E+10	26587.4	429344073404853	160159271	4258220242968	1.14E+19	0.009911	0.000007	107037.5	0.709		100.89	-0.067
			7.45	28.00	5.570	101.08	0.08	0.009894	0.000008	1	1.63E+10	30851.1	503136739080547	161349178	4977796298242	1.55E+19	0.009908	-0.000015	-243054.4	3.622		100.92	0.152
			7.48	30.00	5.715	101.44	0.08	0.009858	0.000008	1	1.65E+10	35430.6	586295304183607	163120544	5779454793665	2.08E+19	0.009905	-0.000048	-789679.9	37.685		100.96	0.489
			sums >>>>>>>					29	4.30E+11		3268256791282030	4268939167	32403050053935	7.19E+19			6.13E-07	135.372
Derived parameters		esd					Chi^2(/n-2) >>>>>		5.014
Denom	2.02E+31					esd of fit >>>>>>		2.239
a	0.009929	0.000004					Ci'measured vs Log(f) (with expectation value)
C1' /pF	100.710	0.043
b	-6.82E-10	3.26E-10
L1 /nH	-0.682	0.326

KT	11.98	< Transformer const. = ( N R0 / k Ri) + (1/N) -1
L2 / nH	900.00	20.00
L1(true)	74.43	1.7	< corrected for L2

				C1'm obs-calc vs Log(f)

## Sheet3
Stray capacitance of voltage sampling network				Predicted C1 = C1' + C1s = C2 [( (1+Cx/C2) N R0 / k Ri ) + (1/N) -1]
		esd	Deriv dC1/dpar	(der*esd)^2
N =	12
R0 =	49.89	0.06	2.74	0.03
C2+Cx /pF =	10.9	0.2	11.63	5.41
Cx /pF	0.3	0.05
C2 /pF	10.6
1+Cx/C2	1.03
Ri =	49.77	0.06	-2.75	0.03
k =	0.959	0.01	-142.57	2.03
				7.49	<<< sum
KT	11.982	< Transformer const. = [ (1+Cx/C2) N R0 / k Ri ] + (1/N) -1

C1' =	100.710	1.008
C1 pred =	127.004	2.737
C1s =	26.294	2.917

C1 prediction for bridge evaluation (see sheet 5)
R0 =	50
C1 pred =	127.031

Self capacitance of transformer secondary				Ci' = Ci apparent + ( Csh * Ro / Ri ) + [ L2 / ( N Ro Rv ) ] + ( LH / Ri^2 ) - Cis - [ L0 / ( R0 Ri ) ] + (C0 Ro / Ri) - [(N-1)Cx R0/Ri] -Cx C2/(N(Cx+C2)
parameter		esd +/-
Csh / pF	0.00	0.00	< shield protrusion capacitance. Set to 0 for gen. side earth.
L2 / nH	900.00	20.00	< inductance of upper voltage sampling arm
L2 fit	1		< =1 when L2 is included in fit on sheet 1, 0 otherwise.
RV0	1531.42		< optimal Rv
LH / nH	15.00	2.00	< inductance of sec. load resistor
Cis / pF	0.40	0.10	< strays accross sec winding
L0 / nH	13.00	3.00	< inductance of pri. transmission line
C0 / pF	3.50	0.50	< capacitance of pri. transmission line
Cneut /pF	0.00	0.00

Series for Ci'		esd +/-	Variance
Ci apparent	0.47	0.09	0.008	< from fit on sheet 1
Corr for Csh	0.00	0.00	0.000
Corr for Cx	-3.33	0.55	0.304
Corr for L2	0.00	0.02	0.000	< inclusion depends on whether L2 is included in fit on sheet 1.
Corr for LH	6.06	0.81	0.652
Corr for Cis	-0.40	0.10	0.010
Corr for L0	-5.24	1.21	1.460
Corr for C0	3.51	0.50	0.251
Ci' / pF	1.07	1.64	2.69	Only valid if parameters have been filled in correctly!

lw / mm	228.00	< length of secondary winding wire
Ci' pred / pF	7.64	< predicted Ci' if phase velocity = c

## Sheet4
Data reduction for 8-48 reference capacitor with 82pF in parallel. Data for dial readings 0 and 18 are excluded from fit (end effects).
		Dial	C /pF	esd C	fit	w	wy	wx	wx^2	wxy	y calc	obs-calc	we	we^2
		x	y					C calc	e
		0	89.0	0.20	0	0.000	0.00	0.00	0.00	0.00	86.92	2.08	0.00	0.000
		1	90.0	0.20	1	25.000	2250.00	25.00	25.00	2250.00	89.46	0.54	13.48	7.269
		2	92.0	0.20	1	25.000	2300.00	50.00	100.00	4600.00	92.00	0.00	-0.06	0.000
		3	94.5	0.20	1	25.000	2362.50	75.00	225.00	7087.50	94.54	-0.04	-1.10	0.049
		4	97.0	0.20	1	25.000	2425.00	100.00	400.00	9700.00	97.09	-0.09	-2.14	0.184
		5	99.5	0.20	1	25.000	2487.50	125.00	625.00	12437.50	99.63	-0.13	-3.19	0.406
		6	102.0	0.20	1	25.000	2550.00	150.00	900.00	15300.00	102.17	-0.17	-4.23	0.715
		7	104.5	0.20	1	25.000	2612.50	175.00	1225.00	18287.50	104.71	-0.21	-5.27	1.111
		8	107.0	0.20	1	25.000	2675.00	200.00	1600.00	21400.00	107.25	-0.25	-6.31	1.593
		9	110.0	0.20	1	25.000	2750.00	225.00	2025.00	24750.00	109.79	0.21	5.15	1.060
		10	112.0	0.20	1	25.000	2800.00	250.00	2500.00	28000.00	112.34	-0.34	-8.39	2.819
		11	115.0	0.20	1	25.000	2875.00	275.00	3025.00	31625.00	114.88	0.12	3.06	0.375
		12	117.5	0.20	1	25.000	2937.50	300.00	3600.00	35250.00	117.42	0.08	2.02	0.164
		13	120.0	0.20	1	25.000	3000.00	325.00	4225.00	39000.00	119.96	0.04	0.98	0.038
		14	122.5	0.20	1	25.000	3062.50	350.00	4900.00	42875.00	122.50	0.00	-0.06	0.000
		15	125.0	0.20	1	25.000	3125.00	375.00	5625.00	46875.00	125.04	-0.04	-1.10	0.049
		16	128.0	0.20	1	25.000	3200.00	400.00	6400.00	51200.00	127.59	0.41	10.36	4.289
		17	130.0	0.20	1	25.000	3250.00	425.00	7225.00	55250.00	130.13	-0.13	-3.19	0.406
		18	131.0	0.20	0	0.000	0.00	0.00	0.00	0.00	132.67	-1.67	0.00	0.000
		sums >>>>>>>>>			17	425.000	46662.50	3825.00	44625.00	445887.50			-1.49E-11	20.527
D	4335000
a	86.919118		Capacitance vs dial reading
esd a	0.014087
b	2.541667
esd b	0.000134
Chi^2/n-2)	1.368464
esd fit	1.169814

## Sheet5
Bridge performance evaluation				FT50-61, 12 turns 0.9mm wire + 2 turn quadrature winding with 250 Ohm pot. Lsec = 9.25uH @ 1.5915MHz. 40pF variable ref cap with padding. No shield. L-bal coil.
N	12
R0	50
C2 /pF	10.9
Li /uH	9.993
Rv0	1528.05	<<< Guide value (estimated value for zero phase error at 0Hz)
Rvset	1528.00	<<< Selected value
C1s /pF	26.294
L1 /nH	-0.682
C1 pred	127.031								| << - - show hidden columns for details of fit - - >> |
Cv set 0	137.931	<<< Guide value (est. value for zero magnitude error at 0Hz)
C1 set	127.300
Cv set	138.200	<<< Selected value
			x						y				| - - - Determining the gradient of Phi vs f: y=phi, x=f
			f	Rv	C1'm	C1'0	Cv	Xcv	Xcvset	U	Rbal	Xbal	|Zbal|	(|Zbal|/R0)-1	Phi				esd phi	fit	weight					Phi calc	obs-calc
		Log(f)	/MHz	/Ohms	/pF	/pF	/pF				Ohms	Ohms	Ohms	Magn. err.	/deg.					w	wx	wy	wxy	wx^2	Y calc	e	we	we^2
		6.20	1.6	1530.7	100.87	100.87	138.07	-724.3	-723.6	3.50E-07	50.056	-0.016	50.056	0.0011	-0.018				0.0500	1	400.00	636.6	-7.3	-11.7	1013.2	0.001	-0.019	-7.675	0.147281
		6.23	1.7	1530.0	100.87	100.87	138.07	-678.1	-677.4	3.58E-07	50.051	-0.007	50.051	0.0010	-0.008				0.0500	1	400.00	680.0	-3.1	-5.3	1156.0	0.001	-0.009	-3.486	0.030374
		6.26	1.8	1529.7	100.87	100.87	138.07	-640.4	-639.8	3.64E-07	50.049	-0.003	50.049	0.0010	-0.004				0.0500	1	400.00	720.0	-1.4	-2.5	1296.0	0.001	-0.004	-1.765	0.007792
		6.28	1.9	1529.7	100.87	100.87	138.07	-606.7	-606.1	3.70E-07	50.049	-0.003	50.049	0.0010	-0.003				0.0500	1	400.00	760.0	-1.4	-2.6	1444.0	0.001	-0.004	-1.713	0.007337
Calibration freq. >>>		6.30	2.0	1529.7	100.87	100.87	138.07	-576.4	-575.8	3.75E-07	50.049	-0.003	50.049	0.0010	-0.003				0.0500	1	400.00	800.0	-1.3	-2.6	1600.0	0.001	-0.004	-1.664	0.006921
		6.34	2.2	1528.3	100.85	100.85	138.04	-524.1	-523.5	3.84E-07	50.052	0.014	50.052	0.0010	0.016				0.0500	1	400.00	880.0	6.5	14.3	1936.0	0.001	0.015	6.156	0.094745
		6.38	2.4	1528.0	100.87	100.87	138.07	-480.3	-479.8	3.90E-07	50.044	0.013	50.044	0.0009	0.015				0.0500	1	400.00	960.0	6.2	14.8	2304.0	0.001	0.014	5.793	0.083890
		6.41	2.6	1532.1	100.87	100.87	138.07	-443.4	-442.9	3.95E-07	50.055	-0.023	50.055	0.0011	-0.027				0.0500	1	400.00	1040.0	-10.7	-27.7	2704.0	0.001	-0.028	-11.023	0.303772
		6.45	2.8	1530.9	100.86	100.86	138.05	-411.7	-411.3	3.99E-07	50.055	-0.011	50.055	0.0011	-0.012				0.0500	1	400.00	1120.0	-5.0	-13.9	3136.0	0.001	-0.013	-5.340	0.071290
		6.48	3.0	1529.8	100.86	100.86	138.05	-384.3	-383.9	4.03E-07	50.053	-0.002	50.053	0.0011	-0.002				0.0500	1	400.00	1200.0	-0.8	-2.4	3600.0	0.001	-0.003	-1.161	0.003368
		6.54	3.5	1527.6	100.83	100.84	138.03	-329.4	-329.0	4.10E-07	50.058	0.015	50.058	0.0012	0.017				0.0500	1	400.00	1400.0	7.0	24.4	4900.0	0.001	0.016	6.596	0.108765
		6.60	4.0	1526.3	100.80	100.80	137.99	-288.3	-287.9	4.15E-07	50.071	0.024	50.071	0.0014	0.027				0.0500	1	400.00	1600.0	10.9	43.5	6400.0	0.001	0.026	10.494	0.275290
		6.65	4.5	1524.5	100.77	100.77	137.97	-256.4	-255.9	4.18E-07	50.079	0.032	50.079	0.0016	0.037				0.0500	1	400.00	1800.0	14.8	66.8	8100.0	0.001	0.036	14.459	0.522627
		6.70	5.0	1523.3	100.73	100.73	137.93	-230.8	-230.3	4.21E-07	50.093	0.037	50.093	0.0019	0.043				0.0500	1	400.00	2000.0	17.0	85.2	10000.0	0.001	0.042	16.646	0.692757
		6.78	6.0	1521.4	100.64	100.64	137.84	-192.4	-191.9	4.24E-07	50.126	0.043	50.126	0.0025	0.049				0.0500	1	400.00	2400.0	19.7	118.2	14400.0	0.001	0.048	19.288	0.930057
		6.85	7.0	1526.1	100.56	100.56	137.75	-165.1	-164.5	4.25E-07	50.161	0.024	50.161	0.0032	0.027				0.0500	1	400.00	2800.0	11.0	76.8	19600.0	0.001	0.026	10.550	0.278230
		6.90	8.0	1533.2	100.62	100.62	137.81	-144.4	-144.0	4.24E-07	50.141	-0.003	50.141	0.0028	-0.003				0.0500	1	400.00	3200.0	-1.3	-10.5	25600.0	0.001	-0.004	-1.750	0.007655
		6.95	9.0	1536.0	100.59	100.59	137.79	-128.3	-128.0	4.24E-07	50.150	-0.009	50.150	0.0030	-0.011				0.0500	1	400.00	3600.0	-4.3	-38.7	32400.0	0.001	-0.012	-4.740	0.056166
		7.00	10.0	1540.7	100.59	100.59	137.79	-115.5	-115.2	4.24E-07	50.151	-0.020	50.151	0.0030	-0.023				0.0500	1	400.00	4000.0	-9.1	-91.0	40000.0	0.001	-0.024	-9.556	0.228305
		7.08	12.0	1546.1	100.57	100.57	137.76	-96.3	-96.0	4.23E-07	50.160	-0.027	50.160	0.0032	-0.031				0.0500	1	400.00	4800.0	-12.3	-147.8	57600.0	0.001	-0.032	-12.798	0.409461
		7.15	14.0	1554.6	100.58	100.58	137.78	-82.5	-82.3	4.21E-07	50.156	-0.038	50.156	0.0031	-0.043				0.0500	1	400.00	5600.0	-17.3	-242.9	78400.0	0.001	-0.045	-17.856	0.797120
		7.20	16.0	1560.7	100.58	100.58	137.78	-72.2	-72.0	4.20E-07	50.156	-0.042	50.156	0.0031	-0.048				0.0500	1	400.00	6400.0	-19.3	-309.4	102400.0	0.001	-0.050	-19.875	0.987534
		7.26	18.0	1546.1	100.56	100.56	137.75	-64.2	-64.0	4.24E-07	50.164	-0.018	50.164	0.0033	-0.020				0.0500	1	400.00	7200.0	-8.1	-146.5	129600.0	0.001	-0.022	-8.701	0.189259
		7.30	20.0	1548.6	100.57	100.57	137.76	-57.8	-57.6	4.23E-07	50.159	-0.019	50.159	0.0032	-0.022				0.0500	1	400.00	8000.0	-8.8	-175.7	160000.0	0.001	-0.023	-9.374	0.219692
		7.34	22.0	1546.9	100.63	100.63	137.83	-52.5	-52.3	4.24E-07	50.136	-0.016	50.136	0.0027	-0.019				0.0500	1	400.00	8800.0	-7.5	-164.9	193600.0	0.002	-0.020	-8.108	0.164366
Cal freq >>		7.38	24.0	1525.6	100.69	100.70	137.89	-48.1	-48.0	4.30E-07	50.112	0.006	50.112	0.0022	0.007				0.0500	1	400.00	9600.0	2.7	65.7	230400.0	0.002	0.005	2.099	0.011013
		7.41	26.0	1514.9	100.83	100.83	138.02	-44.4	-44.3	4.32E-07	50.064	0.014	50.064	0.0013	0.016				0.0500	1	400.00	10400.0	6.6	171.4	270400.0	0.002	0.015	5.926	0.087807
		7.45	28.0	1500.0	101.08	101.08	138.27	-41.1	-41.1	4.36E-07	49.974	0.024	49.974	-0.0005	0.028				0.0500	1	400.00	11200.0	11.2	312.8	313600.0	0.002	0.026	10.482	0.274675
		7.48	30.0	1474.9	101.44	101.45	138.64	-38.3	-38.4	4.42E-07	49.840	0.041	49.840	-0.0032	0.047				0.0500	1	400.00	12000.0	18.8	564.4	360000.0	0.002	0.045	18.098	0.818823
									Sums >>>>		29	1.16E+04	115596.6	13.3	162.1	2077589.2		0.00E+00	0.00E+00	7.816371
									Parameter		esd					Chi^2/n-2 >>>>>>>>>>>			0.289495
phase error vs. f									D	1.07E+10					esd of an obs of unit weight >>>			0.538
									a	0.000822	0.007484	<<< phase intercept at 0Hz
									b	0.000032	0.000559	<<< gradient [degrees / MHz]
									tp	-0.0001	0.001553	ns <<< time delay

									Phase error, obs-calc

Magnitude error vs f