## Sheet1
FT50-61, 12 turns 0.9mm wire + 2 turn quadrature winding with 250 Ohm pot + 120 Ohms. Gen side shield earth. Lsec = 9.25uH @ 1.5915MHz. L-bal coil.					D. W. Knight. July 2007.			Version	1.01
		esd +/-					Show columns N-Q to see deatails of fitting procedure.					See sheet 2 for C1'm vs frequency
N =	12.00					Li' = N R0 C2 Rv
R0 =	49.89	0.06					from fit	from fit
C2	10.30	0.20					Ci =	Li =
L2 / nH	580.00	20.00					Li computed from Rv
L2 fit	1	< 1 to include L2 in fit, 0 to exclude it.
1/(N R0 C2)	1.62E+08					0.068	9.483
Rq	92.04	< DC resistance of the quadrature network.
			Log (f)	f	x	Rv + Rq	Rv	esd Rv	Li'	Corr(L2)	y	esd y	fit	weight	w x	w x^2	w y	w x y	y calc	obs-calc	w e	w e^2	Li' calc	Li'					Computed	Computed
				/MHz	-(2 pi f)^2		Ohms	Ohms	/ uH					w					e			/ uH	obs-calc					Li /uH	Ci /pF
			6.20	1.59	-100.000	1628.2	1536.2	15.36	9.4726	-63.1	105504.76	1055.7	1	8.97E-07	-0.000089730	0.008973	0.094669	-9.466931013	105448.48	56.28	0.000051	0.00	9.4776	-0.0051					9.472	-1.126
			6.23	1.70	-114.093	1628.7	1536.7	7.68	9.4757	-71.9	105461.54	527.7	1	3.59E-06	-0.000409767	0.046751	0.378768	-43.214671955	105447.52	14.03	0.000050	0.00	9.4769	-0.0013					9.475	-0.686
			6.26	1.80	-127.910	1629.3	1537.3	0.92	9.4794	-80.6	105411.67	63.3	1	2.50E-04	-0.031927201	4.083811	26.311452	-3365.499739832	105446.58	-34.91	-0.008713	0.30	9.4762	0.0031					9.479	-0.289
			6.28	1.90	-142.517	1628.6	1536.6	0.92	9.4750	-89.9	105450.49	63.3	1	2.49E-04	-0.035540819	5.065174	26.297174	-3747.796673294	105445.59	4.89	0.001221	0.01	9.4755	-0.0004					9.474	-0.597
Calibration freq. >>>			6.30	2.00	-157.914	1629.0	1537.0	0.92	9.4775	-99.5	105413.34	63.3	1	2.50E-04	-0.039400915	6.221943	26.301598	-4153.381942900	105444.55	-31.21	-0.007787	0.24	9.4747	0.0028					9.477	-0.365
			6.34	2.20	-191.076	1628.3	1536.3	0.92	9.4732	-120.5	105440.46	63.3	1	2.49E-04	-0.047631690	9.101251	26.284406	-5022.307171167	105442.30	-1.84	-0.000459	0.00	9.4730	0.0002					9.472	-0.553
			6.38	2.40	-227.396	1627.6	1535.6	0.92	9.4689	-143.5	105465.61	63.4	1	2.49E-04	-0.056634002	12.878328	26.266723	-5972.939435433	105439.84	25.77	0.006419	0.17	9.4712	-0.0023					9.468	-0.676
			6.41	2.60	-266.874	1628.2	1536.2	0.92	9.4726	-168.3	105399.52	63.3	1	2.49E-04	-0.066518246	17.751997	26.270780	-7010.990976363	105437.16	-37.64	-0.009382	0.35	9.4692	0.0034					9.471	-0.422
			6.45	2.80	-309.511	1626.7	1534.7	0.92	9.4633	-195.4	105475.62	63.4	1	2.49E-04	-0.076994835	23.830733	26.238432	-8121.078071219	105434.27	41.35	0.010288	0.43	9.4670	-0.0037					9.461	-0.697
			6.48	3.00	-355.306	1626.9	1534.9	0.92	9.4646	-224.3	105432.97	63.4	1	2.49E-04	-0.088409968	31.412571	26.234659	-9321.325557514	105431.16	1.81	0.000451	0.00	9.4647	-0.0002					9.462	-0.568
			6.54	3.50	-483.611	1624.9	1532.9	0.92	9.4522	-305.7	105489.44	63.5	1	2.48E-04	-0.120022386	58.044100	26.180349	-12661.094754777	105422.46	66.99	0.016624	1.11	9.4582	-0.0060					9.449	-0.703
			6.60	4.00	-631.655	1624.9	1532.9	0.92	9.4522	-399.2	105395.88	63.5	1	2.48E-04	-0.156763933	99.020672	26.157128	-16522.272188987	105412.42	-16.54	-0.004105	0.07	9.4508	0.0015					9.448	-0.538
			6.65	4.50	-799.438	1623.0	1531.0	0.92	9.4405	-505.9	105420.50	63.6	1	2.48E-04	-0.197912808	158.219011	26.098421	-20864.068091507	105401.04	19.47	0.004819	0.09	9.4423	-0.0017					9.436	-0.589
			6.70	5.00	-986.960	1622.3	1530.3	0.92	9.4362	-624.8	105350.01	63.6	1	2.47E-04	-0.244113415	240.930284	26.057124	-25717.350412833	105388.32	-38.31	-0.009476	0.36	9.4328	0.0034					9.430	-0.526
			6.78	6.00	-1421.223	1621.9	1529.9	0.92	9.4337	-900.0	105102.55	63.6	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	105358.87	-256.31	0.000000	0.00	9.4110	0.0228					9.425	-0.385
			6.85	7.00	-1934.442	1624.0	1532.0	0.92	9.4467	-1223.3	104633.92	63.5	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	105324.06	-690.14	0.000000	0.00	9.3855	0.0612					9.435	-0.208
			6.90	8.00	-2526.619	1631.9	1539.9	0.92	9.4954	-1589.6	103724.55	63.2	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	105283.90	-1559.35	0.000000	0.00	9.3569	0.1385					9.480	0.056
			6.95	9.00	-3197.752	1637.6	1545.6	0.93	9.5305	-2004.4	102921.33	63.0	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	105238.38	-2317.05	0.000000	0.00	9.3246	0.2059					9.511	0.166	Ci computed from Rv
			7.00	10.00	-3947.842	1642.5	1550.5	0.93	9.5608	-2466.8	102127.37	62.8	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	105187.51	-3060.13	0.000000	0.00	9.2890	0.2718					9.536	0.218
			7.08	12.00	-5684.892	1640.0	1548.0	0.93	9.5453	-3557.9	101205.17	62.9	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	105069.70	-3864.53	0.000000	0.00	9.2058	0.3396					9.510	0.122
			7.15	14.00	-7737.770	1633.9	1541.9	0.93	9.5077	-4861.9	100315.68	63.1	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	104930.47	-4614.79	0.000000	0.00	9.1081	0.3996					9.461	0.036
			7.20	16.00	-10106.475	1670.9	1578.9	0.95	9.7359	-6201.4	96511.37	61.6	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	104769.82	-8258.45	0.000000	0.00	9.0113	0.7245					9.671	0.271
			7.26	18.00	-12791.007	1640.8	1548.8	0.93	9.5503	-8001.2	96707.80	62.8	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	104587.75	-7879.95	0.000000	0.00	8.8819	0.6684					9.472	0.058
			7.30	20.00	-15791.367	1622.6	1530.6	0.92	9.4381	-9995.5	95958.62	63.6	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	104384.26	-8425.64	0.000000	0.00	8.7428	0.6952					9.344	-0.032
Cal. freq. >>>			7.34	22.00	-19107.554	1640.1	1548.1	0.93	9.5460	-11957.8	92798.54	62.9	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	104159.35	-11360.81	0.000000	0.00	8.6120	0.9340					9.429	0.037
			7.38	24.00	-22739.569	1578.8	1486.8	0.89	9.1680	-14817.5	94258.00	65.4	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	103913.02	-9655.03	0.000000	0.00	8.4224	0.7455					9.040	-0.159
			7.41	26.00	-26687.410	1558.6	1466.6	0.88	9.0434	-17629.5	92948.37	66.3	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	103645.28	-10696.91	0.000000	0.00	8.2457	0.7977					8.898	-0.192
			7.45	28.00	-30951.079	1527.6	1435.6	0.86	8.8522	-20887.6	92078.16	67.8	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	103356.11	-11277.95	0.000000	0.00	8.0487	0.8035					8.691	-0.243
			7.48	30.00	-35530.576	1501.6	1409.6	0.85	8.6919	-24420.4	90629.07	69.0	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	103045.52	-12416.44	0.000000	0.00	7.8452	0.8467					8.514	-0.270
Derived parameters		esd +/-	sums >>>>>>>>>>>>					14	2.99E-03	-1.162369715	666.615598	315.171685	-122532.786618795			1.31E-13	3.14
D	6.42E-01
a	105455.26	16.49
Li / uH	9.482695	0.0015		y obs - y calc
b = Ci / pF	0.067822	0.0349
chi^2/(n-2)	0.262
esd of fit	0.51
Rv0	1537.8
				Incremental error analysis
X-over /MHz	198.46					R0	49.83	49.89	49.95	err contrib.
				Li	8.994662	9.482695	9.016323	0.010831
				Ci	6.173432	0.067822	6.158601	-0.007416

				C2	10.1	10.3	10.5
				Li	8.828914	9.482695	9.182070	0.176578
				Ci	6.289328	0.067822	6.047431	-0.120949

				Accuracy
				esd Li	0.176916
				esd Ci	0.126106

				Li' vs Log(f)

## Sheet2
Variation of C1'm (nominal capacitance, measured @ 1.5915MHz) with frequency					To see details of fitting calculation show columns N-P
			Log(f)	f	Ci'm	esd C1'm	y	esd y	fit	weight	x	wx	wy	wxy	wx^2	y calc	e	we	we^2		<Ci'm>	obs-calc
				/MHz	/pF		1/C1'm			w					obs-calc
			6.20	1.59	99.98	0.08	0.010002	0.000008	1	1.56E+10	0.0	1116495	156156269	11167.18	7.98E+01	0.010003	-0.000001	-22997.4	0.034		99.97	0.015
			6.23	1.70	99.98	0.08	0.010002	0.000008	1	1.56E+10	14.1	220021189951	156156269	2200652030	3.10E+12	0.010003	-0.000001	-23005.0	0.034		99.97	0.015
			6.26	1.80	99.98	0.08	0.010002	0.000008	1	1.56E+10	27.9	435746119956	156156269	4358332866	1.22E+13	0.010003	-0.000001	-23012.5	0.034		99.97	0.015
			6.28	1.90	99.98	0.08	0.010002	0.000008	1	1.56E+10	42.5	663798188819	156156269	6639309750	2.82E+13	0.010003	-0.000001	-23020.4	0.034		99.97	0.015
			6.30	2.00	99.98	0.08	0.010002	0.000008	1	1.56E+10	57.9	904177396539	156156269	9043582682	5.24E+13	0.010003	-0.000001	-23028.7	0.034		99.97	0.015
			6.34	2.20	99.98	0.08	0.010002	0.000008	1	1.56E+10	91.1	1421917228552	156156269	14222016689	1.30E+14	0.010003	-0.000001	-23046.6	0.034		99.97	0.015
			6.38	2.40	99.97	0.08	0.010003	0.000008	1	1.56E+10	127.4	1988169989977	156109417	19887666200	2.53E+14	0.010003	0.000000	-7442.9	0.004		99.97	0.005
			6.41	2.60	99.97	0.08	0.010003	0.000008	1	1.56E+10	166.9	2604280377748	156109417	26050618963	4.35E+14	0.010003	0.000000	-7464.3	0.004		99.97	0.005
			6.45	2.80	99.97	0.08	0.010003	0.000008	1	1.56E+10	209.5	3269679596540	156109417	32706607948	6.85E+14	0.010003	0.000000	-7487.3	0.004		99.97	0.005
			6.48	3.00	99.97	0.08	0.010003	0.000008	1	1.56E+10	255.3	3984367646354	156109417	39855633153	1.02E+15	0.010003	0.000000	-7512.0	0.004		99.97	0.005
			6.54	3.50	99.97	0.08	0.010003	0.000008	1	1.56E+10	383.6	5986726406608	156109417	59885229635	2.30E+15	0.010003	0.000000	-7581.3	0.004		99.97	0.005
			6.60	4.00	99.95	0.08	0.010005	0.000008	1	1.56E+10	531.7	8290502648791	156015742	82946499738	4.41E+15	0.010003	0.000002	23557.4	0.036		99.97	-0.015
			6.65	4.50	99.95	0.08	0.010005	0.000008	1	1.56E+10	699.4	10906877021602	156015742	109123331882	7.63E+15	0.010003	0.000002	23466.9	0.035		99.97	-0.015
			6.70	5.00	99.95	0.08	0.010005	0.000008	1	1.56E+10	887.0	13831060144156	156015742	138379791337	1.23E+16	0.010004	0.000001	23365.8	0.035		99.96	-0.015
			6.78	6.00	99.95	0.08	0.010005	0.000008	1	1.56E+10	1321.2	20602852638490	156015742	206131592181	2.72E+16	0.010004	0.000001	23131.5	0.034		99.96	-0.015
			6.85	7.00	99.95	0.08	0.010005	0.000008	1	1.56E+10	1834.4	28605880131795	156015742	286201902269	5.25E+16	0.010004	0.000001	22854.6	0.033		99.96	-0.015
			6.90	8.00	99.95	0.08	0.010005	0.000008	1	1.56E+10	2426.6	37840142624070	156015742	378590721601	9.18E+16	0.010004	0.000001	22535.2	0.033		99.96	-0.014
			6.95	9.00	99.95	0.08	0.010005	0.000008	1	1.56E+10	3097.8	48305640115314	156015742	483298050178	1.50E+17	0.010004	0.000001	22173.1	0.032		99.96	-0.014
			7.00	10.00	99.95	0.08	0.010005	0.000008	1	1.56E+10	3847.8	60002372605529	156015742	600323887999	2.31E+17	0.010004	0.000001	21768.4	0.030		99.96	-0.014
			7.08	12.00	99.95	0.08	0.010005	0.000008	1	1.56E+10	5584.9	87089542582868	156015742	871331091374	4.86E+17	0.010004	0.000001	20831.4	0.028		99.96	-0.013
			7.15	14.00	99.95	0.08	0.010005	0.000008	1	1.56E+10	7637.8	119101652556086	156015742	1191612331727	9.10E+17	0.010004	0.000001	19723.9	0.025		99.96	-0.013
			7.20	16.00	99.95	0.08	0.010005	0.000008	1	1.56E+10	10006.5	156038702525185	156015742	1561167609056	1.56E+18	0.010004	0.000001	18446.0	0.022		99.96	-0.012
			7.26	18.00	99.98	0.08	0.010002	0.000008	1	1.56E+10	12691.0	198138399115694	156156269	1981780347226	2.51E+18	0.010004	-0.000002	-29852.1	0.057		99.96	0.019
			7.30	20.00	99.98	0.08	0.010002	0.000008	1	1.56E+10	15691.4	244981526773993	156156269	2450305328806	3.84E+18	0.010004	-0.000002	-31472.7	0.063		99.96	0.020
			7.34	22.00	99.98	0.08	0.010002	0.000008	1	1.56E+10	19007.6	296755509975271	156156269	2968148729499	5.64E+18	0.010004	-0.000002	-33263.8	0.071		99.96	0.021
			7.38	24.00	99.95	0.08	0.010005	0.000008	1	1.56E+10	22639.6	353036302360377	156015742	3532129088148	7.99E+18	0.010004	0.000001	11630.8	0.009		99.96	-0.007
			7.41	26.00	99.95	0.08	0.010005	0.000008	1	1.56E+10	26587.4	414598052308874	156015742	4148054550364	1.10E+19	0.010004	0.000001	9501.1	0.006		99.96	-0.006
			7.45	28.00	99.95	0.08	0.010005	0.000008	1	1.56E+10	30851.1	481084742253251	156015742	4813254049557	1.48E+19	0.010005	0.000000	7200.9	0.003		99.95	-0.005
			7.48	30.00	99.90	0.08	0.010010	0.000008	0	0.00E+00	35430.6	0	0	0	0.00E+00	0.010005	0.000005	0.0	0.000		99.95	-0.053
			sums >>>>>>>					28	4.37E+11		2600688643638880	4370173895	26017628564027	4.94E+19			-2.35E-06	0.774
Derived parameters		esd					Chi^2(/n-2) >>>>>		0.030
Denom	1.48E+31					esd of fit >>>>>>		0.172
a	0.010003	0.000000				Ci'measured vs Log(f) (with expectation value)
C1' /pF	99.965	0.003
b	3.46E-11	2.96E-11
L1 /nH	0.035	0.030

KT	11.63	< Transformer const. = ( N R0 / k Ri) + (1/N) -1
L2 / nH	580.00	20.00
L1(true)	49.92	1.72	< corrected for L2

				C1'm obs-calc vs Log(f)

## Sheet3
Stray capacitance of voltage sampling network				Predicted C1 = C1' + C1s = C2 [( (1+Cx/C2) N R0 / k Ri ) + (1/N) -1]
		esd	Deriv dC1/dpar	(der*esd)^2
N =	12
R0 =	49.89	0.06	2.59	0.02
C2+Cx /pF =	10.3	0.2	11.63	5.41
Cx /pF	0	0
C2 /pF	10.3
1+Cx/C2	1
Ri =	49.77	0.06	-2.6	0.02
k =	0.959	0.0112	-134.72	2.28
				7.73	<<< sum
KT	11.627	< Transformer const. = [ (1+Cx/C2) N R0 / k Ri ] + (1/N) -1

C1' =	99.965	1.000
C1 pred =	119.753	2.781
C1s =	19.788	2.955

C1 prediction for bridge evaluation (see sheet 5)
R0 =	50
C1 pred =	120.038

Self capacitance of transformer secondary				Ci' = Ci apparent + ( Csh * Ro / Ri ) + [ L2 / ( N Ro Rv ) ] + ( LH / Ri^2 ) - Cis - [ L0 / ( R0 Ri ) ] + (C0 Ro / Ri) - [(N-1)Cx R0/Ri] -Cx C2/(N(Cx+C2)
parameter		esd +/-
Csh / pF	0.00	0.00	< shield protrusion capacitance. Set to 0 for gen. side earth.
L2 / nH	580.00	20.00	< inductance of upper voltage sampling arm
L2 fit	1		< =1 when L2 is included in fit on sheet 1, 0 otherwise.
RV0	1537.80		< optimal Rv
LH / nH	15.00	2.00	< inductance of sec. load resistor
Cis / pF	0.40	0.10	< strays accross sec winding
L0 / nH	13.00	3.00	< inductance of pri. transmission line
C0 / pF	3.50	0.50	< capacitance of pri. transmission line
Cneut /pF	0.00	0.00

Series for Ci'		esd +/-	Variance
Ci apparent	0.07	0.03	0.001	< from fit on sheet 1
Corr for Csh	0.00	0.00	0.000
Corr for Cx	0.00	0.00	0.000
Corr for L2	0.00	0.02	0.000	< inclusion depends on whether L2 is included in fit on sheet 1.
Corr for LH	6.06	0.81	0.652
Corr for Cis	-0.40	0.10	0.010
Corr for L0	-5.24	1.21	1.460
Corr for C0	3.51	0.50	0.251
Ci' / pF	4.00	1.54	2.37	Only valid if parameters have been filled in correctly!

## Sheet5
Bridge performance evaluation				FT50-61, 12 turns 0.9mm wire + 2 turn quadrature winding with 250 Ohm pot + 120 Ohms. Gen side shield earth. Lsec = 9.25uH @ 1.5915MHz. L-bal coil.
N	12
R0	50
C2 /pF	10.3
Li /uH	9.483
Rv0	1534.42	<<< Guide value (estimated value for zero phase error at 0Hz)
Rvset	1537.00	<<< Selected value
C1s /pF	19.788
L1 /nH	0.035
C1 pred	120.038								| << - - show hidden columns for details of fit - - >> |
Cv set 0	130.338	<<< Guide value (est. value for zero magnitude error at 0Hz)
C1 set	119.735
Cv set	130.035	<<< Selected value
			x						y				| - - - Determining the gradient of Phi vs f: y=phi, x=f
			f	Rv	C1'm	C1'0	Cv	Xcv	Xcvset	U	Rbal	Xbal	|Zbal|	(|Zbal|/R0)-1	Phi				esd phi	fit	weight					Phi calc	obs-calc
		Log(f)	/MHz	/Ohms	/pF	/pF	/pF				Ohms	Ohms	Ohms	Magn. err.	/deg.					w	wx	wy	wxy	wx^2	Y calc	e	we	we^2
		6.20	1.6	1536.2	99.98	99.98	130.07	-768.8	-769.0	3.39E-07	49.984	0.006	49.984	-0.00031	0.007				0.0500	1	400.00	636.6	2.7	4.3	1013.2	-0.001	0.007	2.987	0.022301
		6.23	1.7	1536.7	99.98	99.98	130.07	-719.8	-720.0	3.47E-07	49.988	-0.001	49.988	-0.00025	-0.001				0.0500	1	400.00	680.0	-0.3	-0.5	1156.0	-0.001	0.000	-0.054	0.000007
		6.26	1.8	1537.3	99.98	99.98	130.07	-679.8	-680.0	3.54E-07	49.991	-0.008	49.991	-0.00018	-0.009				0.0500	1	400.00	720.0	-3.6	-6.5	1296.0	0.000	-0.009	-3.416	0.029164
		6.28	1.9	1536.6	99.98	99.98	130.07	-644.0	-644.2	3.60E-07	49.987	0.001	49.987	-0.00026	0.001				0.0500	1	400.00	760.0	0.3	0.5	1444.0	0.000	0.001	0.373	0.000349
Calibration freq. >>>		6.30	2.0	1537.0	99.98	99.98	130.07	-611.8	-612.0	3.65E-07	49.989	-0.004	49.989	-0.00022	-0.004				0.0500	1	400.00	800.0	-1.8	-3.6	1600.0	0.000	-0.004	-1.749	0.007644
		6.34	2.2	1536.3	99.98	99.98	130.07	-556.2	-556.3	3.74E-07	49.986	0.004	49.986	-0.00028	0.004				0.0500	1	400.00	880.0	1.7	3.7	1936.0	0.000	0.004	1.591	0.006330
		6.38	2.4	1535.6	99.97	99.97	130.06	-509.9	-510.0	3.82E-07	49.987	0.011	49.987	-0.00025	0.013				0.0500	1	400.00	960.0	5.2	12.5	2304.0	0.001	0.013	5.006	0.062649
		6.41	2.6	1536.2	99.97	99.97	130.06	-470.7	-470.7	3.87E-07	49.990	0.005	49.990	-0.00021	0.006				0.0500	1	400.00	1040.0	2.4	6.2	2704.0	0.001	0.005	2.044	0.010441
		6.45	2.8	1534.7	99.97	99.97	130.06	-437.0	-437.1	3.92E-07	49.986	0.018	49.986	-0.00028	0.020				0.0500	1	400.00	1120.0	8.1	22.7	3136.0	0.001	0.019	7.671	0.147107
		6.48	3.0	1534.9	99.97	99.97	130.06	-407.9	-408.0	3.96E-07	49.987	0.015	49.987	-0.00026	0.017				0.0500	1	400.00	1200.0	6.9	20.7	3600.0	0.001	0.016	6.337	0.100386
		6.54	3.5	1532.9	99.97	99.97	130.06	-349.6	-349.7	4.03E-07	49.985	0.027	49.985	-0.00030	0.031				0.0500	1	400.00	1400.0	12.5	43.8	4900.0	0.002	0.029	11.617	0.337383
		6.60	4.0	1532.9	99.95	99.95	130.04	-306.0	-306.0	4.08E-07	49.994	0.026	49.994	-0.00013	0.029				0.0500	1	400.00	1600.0	11.7	47.0	6400.0	0.003	0.026	10.545	0.278010
		6.65	4.5	1531.0	99.95	99.95	130.04	-272.0	-272.0	4.12E-07	49.993	0.034	49.993	-0.00014	0.039				0.0500	1	400.00	1800.0	15.4	69.4	8100.0	0.004	0.035	13.904	0.483314
		6.70	5.0	1530.3	99.95	99.95	130.04	-244.8	-244.8	4.15E-07	49.993	0.034	49.993	-0.00013	0.039				0.0500	1	400.00	2000.0	15.6	78.0	10000.0	0.005	0.034	13.765	0.473667
		6.78	6.0	1529.9	99.95	99.95	130.04	-204.0	-204.0	4.18E-07	49.995	0.030	49.995	-0.00010	0.035				0.0500	1	400.00	2400.0	13.9	83.3	14400.0	0.006	0.029	11.425	0.326307
		6.85	7.0	1532.0	99.95	99.95	130.04	-174.8	-174.8	4.19E-07	49.997	0.018	49.997	-0.00007	0.021				0.0500	1	400.00	2800.0	8.4	58.9	19600.0	0.008	0.013	5.328	0.070956
		6.90	8.0	1539.9	99.95	99.95	130.04	-153.0	-153.0	4.18E-07	50.000	-0.009	50.000	0.00000	-0.011				0.0500	1	400.00	3200.0	-4.2	-34.0	25600.0	0.009	-0.020	-7.953	0.158131
		6.95	9.0	1545.6	99.95	99.95	130.04	-136.0	-136.0	4.18E-07	50.001	-0.024	50.001	0.00002	-0.028				0.0500	1	400.00	3600.0	-11.2	-100.7	32400.0	0.011	-0.039	-15.521	0.602259
		7.00	10.0	1550.5	99.95	99.95	130.04	-122.4	-122.4	4.17E-07	50.002	-0.034	50.002	0.00003	-0.039				0.0500	1	400.00	4000.0	-15.8	-157.9	40000.0	0.012	-0.052	-20.742	1.075585
		7.08	12.0	1548.0	99.95	99.95	130.04	-102.0	-102.0	4.18E-07	50.000	-0.023	50.000	0.00001	-0.027				0.0500	1	400.00	4800.0	-10.8	-129.1	57600.0	0.016	-0.042	-16.962	0.719306
		7.15	14.0	1541.9	99.95	99.95	130.04	-87.4	-87.4	4.21E-07	49.999	-0.009	49.999	-0.00001	-0.010				0.0500	1	400.00	5600.0	-4.1	-57.8	78400.0	0.019	-0.029	-11.583	0.335394
		7.20	16.0	1578.9	99.95	99.95	130.04	-76.5	-76.5	4.11E-07	50.002	-0.066	50.002	0.00004	-0.075				0.0500	1	400.00	6400.0	-30.2	-483.1	102400.0	0.022	-0.097	-38.902	3.783389
		7.26	18.0	1548.8	99.98	99.98	130.07	-68.0	-68.0	4.19E-07	49.988	-0.017	49.988	-0.00024	-0.020				0.0500	1	400.00	7200.0	-7.9	-142.9	129600.0	0.025	-0.045	-17.897	0.800725
		7.30	20.0	1530.6	99.98	99.98	130.07	-61.2	-61.2	4.24E-07	49.987	0.008	49.987	-0.00026	0.009				0.0500	1	400.00	8000.0	3.6	72.0	160000.0	0.028	-0.019	-7.607	0.144678
Cal. freq. >>>		7.34	22.0	1548.1	99.98	99.98	130.07	-55.6	-55.6	4.20E-07	49.988	-0.013	49.988	-0.00024	-0.015				0.0500	1	400.00	8800.0	-6.1	-134.8	193600.0	0.031	-0.046	-18.587	0.863674
		7.38	24.0	1486.8	99.95	99.95	130.04	-51.0	-51.0	4.37E-07	49.997	0.056	49.997	-0.00006	0.064				0.0500	1	400.00	9600.0	25.6	615.6	230400.0	0.034	0.030	11.939	0.356352
		7.41	26.0	1466.6	99.95	99.95	130.04	-47.1	-47.1	4.43E-07	49.996	0.073	49.997	-0.00007	0.084				0.0500	1	400.00	10400.0	33.7	875.3	270400.0	0.037	0.047	18.704	0.874605
		7.45	28.0	1435.6	99.95	99.95	130.04	-43.7	-43.7	4.53E-07	49.996	0.100	49.996	-0.00008	0.115				0.0500	0	0.00	0.0	0.0	0.0	0.0	0.041	0.074	0.000	0.000000
		7.48	30.0	1409.6	99.90	99.90	129.99	-40.8	-40.8	4.61E-07	50.015	0.120	50.015	0.00029	0.138				0.0500	1	400.00	12000.0	55.2	1655.9	360000.0	0.044	0.094	37.737	3.560185
									Sums >>>>		28	1.12E+04	104396.6	126.9	2418.9	1763989.2		7.45E-02	5.37E-14	15.630297
									Parameter		esd					Chi^2/n-2 >>>>>>>>>>>			0.601165
phase error vs. f									D	8.86E+09					esd of an obs of unit weight >>>			0.775
									a	-0.003243	0.010941	<<< phase intercept at 0Hz
									b	0.001563	0.000872	<<< gradient [degrees / MHz]
									tp	-0.0043	0.002422	ns <<< time delay

									Phase error, obs-calc

Magnitude error vs f